## 2B-1-2-2026 G1 Methods
| Cell Preparation and Controls | [12 empty] | Sensitisation Stage |  | First Wash Step |  |  |  |  | Detection Stage | [8 empty] | Anti-T Cell Reagent | [17 empty] | Anti-B-cell Reagents | [17 empty] | Final Wash | [7 empty] | Data Acquisiton | [14 empty] | Data Analysis
Laboratory ID | Cell Preparation Method | Cell Volume/Tube | Cell Concentration | Cell Diluent | Serum Volume/Tube | Pronase Treatment | Pronase Treatment Details | Serum Dilutions | Negative Control | Number of Negative Replicates | Positive Control | Number of Positive Replicates | Number of Crossmatch Test Replicates | Incubation Time (Mins) | Sensitisation Temperature | First Wash Step | Wash Medium | First Wash Volume/Tube | Number of Washes | First Wash Temperature | Manufacturer | Anti-human IgG Reagent Raised In | Anti-human IgG Labelled With | Detection Stage More Details | Detection Stage Dilution | Detection Stage Diluted In | Detection Stage Volume / Tube | Detection Stage Incubation Time (mins) | Detection Stage Incubation Temperature | Anti-T Cell Reagent Manufacturer | Anti-Human IgG Reagent added with Anti-T cell | Is there a wash step before adding Reagent | Anti-T Cell Reagent Wash Step Details | Anti-T Cell Reagent Wash Medium | T-Cell Volume/Tube | Anti-T Cell Reagent Number of Washes | T-Cell Incubation Temp | Is Anti T cell reagent with Anti-B cell reagent | Anti-T Cell Reagent Labelled with | Anti-T Cell Reagent Other Details | Anti-T Cell Reagent CD3 | Anti-T Cell Reagent Diluted | Anti-T Cell Reagent Diluted In | Anti-T Cell Reagent Dilution | Anti-T Cell Reagent Volume/Tube | T-Cell Incubation Time (mins) | Anti-T Cell Reagent Temperature | Anti B-Cell Reagent Manufacturer | Is Anti B Cell Reagent added with Anti human IgG | Is there a wash step before adding Anti B Cell | Anti B-Cell Reagent Wash Step Details | Anti B-Cell Reagent Wash Medium | B-Cell Volume/Tube | Anti B-Cell Reagent Number of Washes | Anti B-Cell Reagent Temperature | Anti B-Cell Reagent Labelled With | Anti B-Cell Reagent Other Details | Anti B-Cell Reagent CD19 | Anti B-Cell Reagent CD20 | Anti B-Cell Reagent Diluted | Anti B-Cell Reagent Diluted In | Anti B-Cell Reagent Dilution | Anti B-Cell Reagent Volume/Tube | B-Cell Incubation Time (mins) | B-Cell Incubation Temp | Final Wash Step Details | Final Wash Medium | Final Wash Volume/Tube (ml) | Final Wash Number of Washes | Final Wash Temperature (C) | Final Wash Resuspension of Cells | Final Wash Other Details | Final Wash Volume/Tube (µl) | Data Acquisition Make of Flow Cytometer | Data Acquisition Live Gated | Data Acquisition Calibration Beads Used | Data Acquisition Used For | Data Acquisition Type of Beads | Total Events Minimum | Total Events Maximum | T-Cells Minimum | T-Cells Maximum | B-Cells Minimum | B-Cells Maximum | Data Acquisition Parameter Assess Peak Performance | Data Acquisition Peak Performance Please Specify | Data Acquisition Scale Used | Data Acquisition Scale Used Please Specify | Data Analysis Parameter used to Assess Difference | Data Analysis Other Details | Data Analysis T Cell Miniumum Value for PCS | Data Analysis B Cell Minimum Value for PCS | Data Analysis Positive From Negative | Data Analysis Distinguish Binding of Human IgG | Comments
Lab-1003 | CELL DENSITY CENTRIFUGATION | 50 | 5x10^6 | PBS | 50 | No |  | No | In- House | 4 | NIBSC | 2 | 4 | 15 | 37 | Automated cell washer | PBS | 2 | 2 | 2-8 | Jackson Immunoresearch | Goat | FITC |  | 5% | CD3-PC5 AND CD19-PE | 5 | 15 | 2-8 | Beckman Coulter | Yes |  |  |  | 5 |  | 2-8 | Yes | Phycoerythrin-cyanine 5 |  | Yes | Yes | CD19-PE AND ANTI-HUMAN IgG -FITC | 47.5% |  | 15 |  | Beckman Coulter | Yes |  |  |  | 5 |  |  | Phycoerythrin |  | Yes | No | Yes | CD3-PC5 AND ANTI-HUMAN IgG-FITC | 47.5% |  | 15 | 2-8 | Automated cell washer | PBS | 2 | 1 | 2-8 | Phosphate Buffered Saline |  | 300 | Beckman Coulter DxFlex | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | DX FLEX QC BEADS | N/A | N/A | 250 | N/A | 100 | 1000 | Median |  | Log |  | Median difference |  | 60 Linear channel shift | 60 Linear channel shift |  | 40 Linear channel shift
Lab-1008 | Density centrifugation | 30 | 1 x 10/6 per ul | Flow diluent (in house) | 20 | No |  | No | NIBSC | 2 | In house pooled serum | 2 | 2 | 20 | 21 | Manual | Flow diluent (in house) | 0.2 | 3 | 21 |  | Rabbit | FITC |  | Neat |  | 4 | 15 | 4 | Dako | Yes |  | Manual | Flow diluent (in house) | 4 | 3 | 4 | Yes | Phycoerythrin |  | Yes | No |  |  | 0.2 | 15 | 21 | Dako | Yes | No | Manual | Flow diluent (in house) | 4 | 3 | 21 | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  | 0.2 | 15 | 4 | Manual | Flow diluent (in house) | 0.5 | 1 | 21 | Other | Flow diluent (in house) | 500 | Beckman Coulter | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | DXFlex Daily QC Fluorospheres |  |  | 5000 |  | 1500 |  | Median |  | Log |  | Median difference | Relative median fluorescence (RMF) | 4 RMF | 4 RMF | Positive control is at least four times that of the negative control (RMF >4.0)
Lab-1019 | Density gradient centrifugation | 50 | 1X7^7 cells/mL | PBS +1% New born calf serum (NCS) | 50 | No |  | No | NIBSC | 3 | Anti-HLA Monoclonal antibody | 2 | 2 | 30 | 37 | Manual | PBS +1% New born calf serum (NCS) | 4 | 3 | 22 | Southern Biotech | Goat | FITC |  | 1:100 | PBS:1% NCS | 20 | 30 | 22 | Beckman-Coulter | Yes | Yes | Manual | PBS +1% New born calf serum (NCS) | 5 | 1 | 22 | Yes | Other | Phycoerythrin | Yes | No |  |  | 4 | 30 | 22 | Beckman-Coulter | Yes | Yes | Manual | PBS +1% New born calf serum (NCS) | 5 | 1 | 22 | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  | 4 | 30 | 22 | Manual | PBS +1% New born calf serum (NCS) | 4 | 1 | 22 |  | PBS +1% New born calf serum (NCS) | 300 | Beckman-Coulter | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | Flow-Check/Flow-Set Beads |  |  | 15990 | 21385 | 1416 | 1910 | Median |  | Log |  | Median difference |  | RMF(Mean test sample Median/Mean Negative Control Median)&gt;1.5 | RMF(Mean test sample Median/Mean Negative Control Median)&gt;5 | T Cell RMF&gt;1.5, B Cell RMF&gt;3 | T Cell RMF&gt;2, B Cell RMF&gt;5
Lab-1043 | Lympholyte DGM | 50 | 1.0 x 10 7/ml | 1% FBS/PBS buffer | 50 | No | N/A | No | NIBSC | 3 | In-house | 2 | 3 | 30 | 37 | Manual | 1% FBS/PBS buffer | 4 | 3 | 22 | Sigma | Goat | FITC |  | 1:5 | 1% FBS/PBS buffer | 4 | 20 | 6 | Dako | No | Yes | Manual | 1% FBS/PBS buffer | 4 | 1 | 6 | Yes | Phycoerythrin |  | Yes | No |  |  | 4 | 20 | 22 | Dako | No | Yes | Manual | 1% FBS/PBS buffer | 4 | 1 | 22 | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  | 4 | 20 | 6 | Manual | 1% FBS/PBS buffer | 4 | 1 | 22 | Other | 1% FBS/PBS buffer | 250 | Beckman Coulter | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check | DxFlex Daily QA Fluorospheres/Flow Check pro & Flow Set pro |  |  | 500 |  | 500 |  | Median |  | Log |  | Other | Relative Mean Fluorescence ratio | >1.5 RMF ratio | >2.0 RMF ratio | >1.5 RMF/>2.0 RMF ratios | >1.5 RMF/>2.0 RMF ratios
Lab-1055 | Miltenyi MACSPrep | 50 | 5million/ML | 1% PBSA | 50 | No |  | No | NIBSC | 3 | NIBSC | 2 | 3 | 15 | 37 | Automated cell washer | 1% PBSA | 2.5 | 3 | RT | Becton Dickinson | Goat | FITC |  | N/A | N/A | 5 | 15 | 2-8 | Becton Dickinson | No | Yes | Automated cell washer | 1% PBSA | 5 | 3 | 2-8 | Yes |  | BV421 | Yes | No |  |  | 2.5 | 15 | RT | Becton Dickinson | No | Yes | Automated cell washer | 1% PBSA | 5 | 3 | RT |  | APC | Yes | No | No |  |  | 2.5 | 15 | 2-8 | Automated cell washer | 1% PBSA | 2.5 | 3 | RT | Other | 1% PBSA | 500 | BD FACSLyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check | BD CS&T Beads | N/A | N/A | N/A | N/A | 500 | N/A | Median |  | Log |  | Other | RMF | Median >1000 | Median >2000 | N/A | T cell RMF >1.2, B cell RMF >1.3
Lab-1079 | Immunomagnetic negative selection | 25 | 4.0x10^6/ml | PBS-BSA-Sodium Azide | 25 | No |  | No | Male AB serum - Life Sciences | 3 | Pooled patient serum | 1 | 3 | 25 | 22 | Manual | PBS-BSA-Sodium Azide | 2 | 2 | 22 | DAKO Agilent | Rabbit | FITC |  |  |  | 10 | 25 | 22 | DAKO - Agilent | Yes | No |  |  | 2.5 |  | 22 | Yes | Phycoerythrin |  | Yes | No |  |  |  | 25 |  | DAKO-Agilent | Yes | No |  |  | 2.5 |  |  | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  |  | 25 | 22 | Manual | PBS-BSA-Sodium azide | 2 | 3 | 22 | Other | PBS-BSA-Sodium azide | 280 | FACS Lyric | No | Yes | Alignment Check, Laser Check, Fluorescence Check | CS&T beads |  |  | 6000 |  | 1250 |  | Median |  | Log |  | Median difference |  | 1.3 RMF | 1.5 RMF
Lab-1108 | Density Gradient Centrifugation | 50 | 5 million per ml | PBS | 50 | No | If enough cells present pre pronase e.g. 10 million per ml. The concentration is 0.5mg/ml. | No | Commercial male Ab serum | x 4 | pooled sera from local positive patients | x2 | x3 | 30 | 37 | Manual | PBS/BSA/Sodium Azide for non-pronased; PBS for pronased | 4 | 3 | 2-8 | Sigma | Goat | FITC |  | Batch specific e.g. 1:50 | PBS | 5 | 30 | 2-8 | Becton Dickinson | Yes | No | Manual | PBS/BSA/Sodium Azide for non-pronased; PBS for pronased | 5 | 1 | 2-8 | Yes | Other | BV421 | Yes | Yes | PBS | e.g 1:1 | 4 | 30 | 2-8 | Becton Dickinson | Yes |  | Manual | PBS/BSA/Sodium Azide for non-pronased; PBS for pronased | 5 | 1 | 2-8 | Other | Allophycocyanin | Yes | No | Yes | yes | e.g. 1:4 | 4 | 30 | 2-8 | Manual | PBS | 4 | 1 | 2-8 | Phosphate Buffered Saline |  | 300ul | Becton Dickinson FACS Lyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | Manufacturer recommendations -Cytometer Setup and Tracking (CS&T) beads |  |  |  |  | 100 | 1000 | Median |  | Log |  |  | 1.5 times the mean of the x4 negative replicate median values | 1.5 times the mean of the x4 negative replicate median values | 1.5 times the mean of the x4 negative replicate median values | 1.5 times the mean of the x4 negative replicate median values | 1.5 times the mean of the x4 negative replicate median values | MU taken into consideration via BSHI publication 2016- Recommendations for the determination of measurement uncertainty for assays performed in the histocompatibility and immunogenetics (H&I) laboratory. This is a type A MU - uncertainty estimates using statistics.
Lab-1118 | Negative selection - BD MACSPrep B/T Separation Beads | 10ul | 10x10^6 cells/ml | Sodium Citrate/2% FCS | 20ul | Yes | 4.7U/ml | No | NIBSC | 2 | In-house | Dilution series - 1:8, 1:16, 1:32, 1:64, 1:128 | 3 | 30 | 20-25 | Automated cell washer | Flow Cell Wash (BD) | 0.2 | 4 | 20-25 | Sigma (Merck) | Goat | FITC |  | 1:80 | PBS | 5 | 10 | 20-25 | BD | Yes |  |  |  | 0.5 |  | 20-25 | Yes | Other | Allophycocyanin (APC) | Yes | No |  |  |  | 10 |  | BD | Yes |  |  |  | 0.5 |  |  | Other | Brilliant Violet 421 (BV421) | Yes | No | No |  |  |  | 10 | 20-25 | Automated cell washer | Flow Cell Wash (BD) | 0.15 | 1 | 20-25 | Other | BD Flow Cell Wash | 200 | BD FACSLyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | BD Cytometer Set-up and Tracking Beads (CST) | N/A | N/A | 200 | N/A | 200 | N/A | Median |  | Other | Relative Ratio (RR) - MCFN Sample / MCFN Negative | Median difference | Relative Ratio (RR) | RR > 1.5 | RR > 2.0 | T-cell RR < 1.5 B-cell RR >2.0
Lab-1128 | DENSITY GRADIENT | 50 |  | PBS | 50 | No |  |  | NIBSC | 2 | IN HOUSE | 1 | 2 | 30 | 37 | Manual | PBS+FBS | 2 | 3 |  |  | Goat | FITC |  | 1:32 | PBS | 50 | 30 | 4 | Beckman | Yes | [5 empty] | 4 | Yes | Phycoerythrin |  | Yes | No |  |  |  | 30 |  | Beckman | Yes | [6 empty] | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  |  | 30 | 4 | Manual | pbs | 2 | 1 |  | Phosphate Buffered Saline |  | 200 | Beckman | Yes | Yes |  |  |  |  | 1000 |  | 300 |  | Median |  | Log |  | Other | log median x | 1.904*av neg control | 2.795*av neg control
Lab-1139 | StemCell Technologies Total Lymphocyte Preparation | 15 | 10.0 x 10^6/ml | PBS+2%FBS | 30 | Yes | 4.7 Units/mL | No | Sigma-Aldrich | 3 | NIBSC | 1 | 3 | 20 | 22 | Manual | PBS+0.1%Azide | 200 | 4 | 22 | Jackson Scientific | Goat | FITC |  | 1 in 400 | PBS | 0.125uL Conjugate in 50uL stain cocktail | 5 minutes | 22C | Becton Dickinson | Yes |  | Manual | PBS 0.1% Azide | 50 | 4 | 22C | Yes | Other | PerCP | Yes | Yes | PBS | 1 in 25 | 2uL neat | 5 minutes | 22C |  | Yes |  | Manual | PBS 0.1% Azide | 50ul | 4 | 22C | Phycoerythrin |  | Yes | No | Yes | PBS | 1 in 50 | 1uL neat | 5 minutes | 22C | Manual | PBS+0.1%Azide | 200 | 2 | 22 | Sheath fluid |  | 100 | BD FACSLyric Flow Cytometer or BD FACSCanto | Yes | Yes | Alignment Check, Laser Check | BD Cytometer Set up and tracking beads | N/A | N/A | N/A | N/A | 200 | N/A | Median |  | Linear |  | Median difference, Other | Median channel shift ratio if acquired on FACSLyric, Median channel shift difference if acquired on FACSCanto | Ratio of 1.595 or greater for Lyric acquistions, a difference of 50 or more MCS for CANTO acquisitions | Ratio of 2.545 of greater, a difference of 100 or more MCS for CANTO acquisitions | As above | As above
Lab-1153 | Lymphoprep | 40 | 2M/ml | McCoys | 10 | No |  | No | Local Ab and NIBSC | 2 | Local and NIBSC | 2 | 2 | 30 | room temp | Manual | PBA | 1ml | 2 | room temp | SIGMA | Goat | FITC |  | 1 in 32 | PBA | 50 | 45 | ice |  | No | Yes | Manual | PBA | 40 | 1 | ICE | Yes | Phycoerythrin |  | Yes | Yes | PBA | 1 in 10 | 1 | 15 | 4 |  | No | Yes | Manual | PBA | 50 | 1 | 4 | Phycoerythrin-cyanine 5 |  | Yes | No | Yes | PBA | 1 in 25 | 1 | 15 | ice | Manual | PBA | 1 | 1 | 4 | Other | PBA | 300 | Becton Dickinson | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | BD CSandT |  |  |  |  | 300 | 1000 | Median |  | Log |  | Median difference |  | 40 channel shift | 40 channel shift
Lab-1174 | Lympholyte | 50 | 5x10^6 per ml | PBS/FBS | 50 | No |  | No | NIBSC | 3 | NIBSC | 2 | 3 | 30 | 37 | Automated cell washer | PBS/FBS |  |  | PBS/FBS at 4-8 | Sigma | Goat | FITC |  | 1:40 | PBS | 50 | 30 | 4 | BD | Yes |  | Automated cell washer | PBS/FBS | 2 |  | 4 | Yes | Other | APC | Yes | No |  |  |  | 30 | 4 | BD | Yes |  | Automated cell washer | PBS/FBS | 2 |  | 4 | Other | BV421 | Yes | No | No |  |  |  | 30 | 4 | Automated cell washer | PBS/FBS |  |  | 4 | Fixative |  | 400 | BD | No | Yes | [10 empty] | Other | Mean +3SD | Other | Mean +3SD
Lab-1176 | Accuspin | 20 | 5 million / ml | PBS/Azide | 20 | No |  | No | NIBSC | 2 | Pooled patient sera | 2 | 2 | 20 | 22 | Automated cell washer | PBS/Azide | 5 | 3 | 22 | Southern Biotech | Goat | FITC |  | 1 in 50 | PBS/Azide | 2 | 10 | 4 | Beckman Coulter | Yes |  |  | N/A | 5 | N/A | 4 | Yes | Other | APC | Yes | Yes | PBS/Azide | 1 in 20 | N/A | 10 | N/A | Beckman Coulter | Yes |  |  | N/A | 5 | N/A | N/A | Other | PC7 | Yes | No | Yes | PBS/Azide | 1 in 20 | N/A | 10 | 4 | Automated cell washer | PBS/Azide | 5 | 3 | 22 | Other | PBS/Azide | 300 | BD FACS Lyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | BD CS&T beads | N/A | N/A | 400 | N/A | 400 | N/A | Median |  | Other | Relative Fluorescence Intensity | Other | RFI | 4.5 RFI | 3.0 RFI | 1.67 RFI for T cells, 1.61 RFI for B cells | 1.67 RFI for T cells, 1.61 RFI for B cells
Lab-1199 | Centrifugation over Density gradient followed by differential washing | 25 | 1 million per ml | 2% FBS/PBS | 50 | No |  | no | BAG | 3 | Invivogen | 3 | 3 | 20 mins | 20-22 | Manual | 2% FBS/PBS | 0.15 | 4 | 20-22 | SIGMA | Goat | FITC |  | 1:3000 | 2% FBS/PBS | 85 | 10 | 20-22 | BD | Yes |  |  |  | 5 |  | 20-22 | Yes | Other | APC | Yes | No |  |  |  | 10 |  | BD | Yes |  |  |  | 1.25 |  |  | Other | BV-421 | Yes | No | No |  |  |  | 10 | 20-22 | Manual | 2% FBS/PBS | 150 | 2 | 20-22 | Other | 2% FBS/PBS | 300 | BD FACS LYRIC | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | CS&T | 50 | 500 | 50 | 500 | 50 | 500 | Median |  | Linear |  | Other | RMF | RMF≥16 | RMF≥16 | RMF<1.0 | RMF≥1.6
Lab-1212 | lymphoprep | 30 | 1000/μl | Flow buffer | 40 | No |  | No | NIBSC | 2 | In house - pooled highly sensitised patient sera | 2 | 2 | 30 minutes | Room temperature | Manual | Flow buffer | 2mL | 2 | Room temperature | Agilent Technologies | Rabbit | FITC |  |  |  | 15uL | 20 minutes | 2-8°C | VHBIO | Yes | No | Manual |  | 11uL |  | 2-8°C | Yes | Phycoerythrin |  | Yes | No |  |  |  | 20 minutes |  | Dako | Yes | No | Manual |  | 5.4uL |  |  | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  |  | 20 minutes | 2-8°C | Manual | Flow buffer | 2mL | 2 | Room temperature | Other | Flow buffer | 0.3mL | BD | Yes | Yes |  |  |  |  | 1,000 |  | 1,000 | 3,000 | Median |  | Log |  | Other | RMF of sample to Negative control |  |  | RMF 1.3 or above is positive for detection of analyte | N/A
Lab-1216 | Density gradient centrifugation using SepMate tubes | 200 | 2.5x106 | PBS | 50 | No |  | yes | Commercial serum from Male AB blood group | 5 | In house pool positive serums | 2 | None | 20 | 37 | Manual | PBS | 100 ul | 3 | room | Invitrogen | Goat | Fab2 fragment | Labelled with FITC | 1/40 | PBS | 50 | 20 | 4 | Beckman Coulter | Yes | No | Manual |  | 50 |  | 4 | Yes | Phycoerythrin |  | Yes | Yes | PBS | 1/40 |  | 20 |  | Beckman Coulter | Yes |  |  |  | 50 |  |  | Other | PC7 | Yes | No | Yes | PBS | 1/40 |  | 20 | 4 | Manual | PBS | 100 ul | 2 | room | Phosphate Buffered Saline |  | 400 | Beckman Coulter DxFlex | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | DxFlex Daily QC fluorospheres & Flow-Set Pro Fluorospheres | 50000 | 100000 | 18000 | 20000 | 3000 | 3000 | Median |  | Log |  | Median difference |  | >1.5 times AB control trimmed mean value | >3 times AB control trimmed mean value | as above for T & B cells | as above for T & B cells
Lab-1238 | Density gradient centrifugation | 50 | 6x10^6 | Tris Buffered Saline | 50 | No | - | No | AB Negative Serum | 3 | In house | 2 | 3 | 30 | 35-39C | Manual | Tris Buffered Saline | 3.5 | 2 | 4-8C | Sigma | Goat | FITC |  | 1/32 | Tris Buffered Saline | 50 | 30 | 4-8C | BD | Yes | No | Manual |  | 5 |  | 4-8C | No | Phycoerythrin |  | Yes | No | - | - |  | 30 |  | BD | Yes | No | Manual |  | 5 |  |  | Phycoerythrin |  | No | Yes | No | - | - |  | 30 | 4-8C | Manual | Tris Buffered Saline | 3.5 | 1 | 4-8C | Other | Tris Buffered Saline | 280 | Beckman Coulter DxFlex | No | Yes | Alignment Check, Laser Check | DxFLEX Daily QC Fluorospheres | - | - | 100 | 3000 | 100 | 3000 | Median |  | Log |  | Median difference |  | AB serum average + 2SDs | AB serum average + 2SDs | AB serum average + 2SDs | AB serum average + 2SDs
Lab-1239 | Ficoll-Paque | 30 | 5x10^6/ml | PBS | 20 | No |  | No | Sigma | 4 | In house | 1 Weak positive, 1 Strong positive | 2 | 30 | 37 | Manual | PBS + 0.1% Sodium Azide | 4 | 3 |  | Sigma-Aldrich | Goat | FITC |  | 1 in 20 | PBS + 0.1% Azide | 10 | 20 | 4 | BD | No | Yes | Manual | PBS +0.1% Azide | 3 | 1 | 4 | Yes | Other | PERCP | Yes | No |  |  | 4 | 20 |  | BD | No | Yes | Manual | PBS + 0.1% Azide | 3 | 1 |  | Other | PE | Yes | No | No |  |  | 4 | 20 | 4 | Manual | PBS + 0.1% Azide | 4 | 1 |  | Other | PBS + 0.1% Azide | 400 | BD | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | CS&T |  |  | 500 |  | 200 |  | Median |  | Log |  | Peak shift |  | 1.3RMF | 1.2RMF
Lab-1264 | density separation | 25 ul (CD3), 50 ul (CD19) | 10x10^6/ml | FACS diluent (in house) | 25 ul (CD3), 50 ul (CD19) | No |  | no | NHSBT | 2 | In house | 2 | 2 | 30 | 22 | Manual | PBS | 2 | 3 | 22 | SIGMA | Goat | FITC |  | 1/140 | FACS Diluent (in house) | 100 | 15 | 4 | AGILENT TECHNOLOGIES (DAKO) | No | Yes | Manual | PBS | 100 | 2 | 4 | No | Phycoerythrin |  | Yes | Yes | FACS diluent (in house) | 1/70 | 2 | 15 | 22 | AGIELNT TECHNOLOGIES (DAKO) | No | Yes | Manual | PBS | 100 | 2 | 22 | Phycoerythrin |  | Yes | No | Yes | FACS DILUENT (IN HOUSE) | 1/40 | 2 | 15 | 4 | Manual | PBS | 2 | 2 | 22 | Other | FACS DILUENT (IN HOUSE) | 200 | DX FLEX (BECKMAN COULTER) | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | QC FLUROSPHERES | 300 | 7500 | 300 | 7500 | 300 | 7500 | Median |  | Other | RMF (RELATIVE MEDIAN FLUORESCENCE) | Other | RELATIVE MEDIAN FLUORESCENCE (RMF) | 10 | 10 | N/A | RMF 1.5

## hiddenSheet (hidden)
vunhst_methodology_scheme2b:6FxwlK/Vqm3xDgAGWbsmXdhKrp+PJtpEDkMRimyftL4SS1HmVc7g8SW9ZLf4W9sQIjuhfBw6E820yAJrksqJ/A==:vunhst_methodology_scheme2bid=%28Do%20Not%20Modify%29%20Methodology_Scheme2B&checksumLogicalName=%28Do%20Not%20Modify%29%20Row%20Checksum&modifiedon=%28Do%20Not%20Modify%29%20Modified%20On&38ac67f9-4cd9-4988-9e75-ada38a03ebb7.vunhst_laboratorycode=Laboratory%20Code%20%28Laboratory%29%20%28Laboratory%29&vunhst_laboratory=Laboratory&vunhst_sample=Sample&vunhst_registration=Registration&createdon=Created%20On&vunhst_antibcellreagentcd19=Anti%20B-Cell%20Reagent%20CD19&vunhst_antibcellreagentcd20=Anti%20B-Cell%20Reagent%20CD20&vunhst_antibcellreagentdiluted=Anti%20B-Cell%20Reagent%20Diluted&vunhst_antibcellreagentdilutedin=Anti%20B-Cell%20Reagent%20Diluted%20In&vunhst_antibcellreagentdilution=Anti%20B-Cell%20Reagent%20Dilution&vunhst_antibcellreagentlabelledwith=Anti%20B-Cell%20Reagent%20Labelled%20With&vunhst_antibcellreagentmanufacturer=Anti%20B-Cell%20Reagent%20Manufacturer&vunhst_antibcellreagentnumberofwashes=Anti%20B-Cell%20Reagent%20Number%20of%20Washes&vunhst_antibcellreagentotherdetails=Anti%20B-Cell%20Reagent%20Other%20Details&vunhst_antibcellreagenttemperature=Anti%20B-Cell%20Reagent%20Temperature&vunhst_antibcellreagentvolumetube=Anti%20B-Cell%20Reagent%20Volume%2fTube&vunhst_antibcellreagentwashmedium=Anti%20B-Cell%20Reagent%20Wash%20Medium&vunhst_antibcellreagentwashstepdetails=Anti%20B-Cell%20Reagent%20Wash%20Step%20Details&vunhst_antihumanigglabelledwith=Anti-human%20IgG%20Labelled%20With&vunhst_antihumaniggreagentaddedwithantitcell=Anti-Human%20IgG%20Reagent%20added%20with%20Anti-T%20cell&vunhst_antihumaniggreagentraisedin=Anti-human%20IgG%20Reagent%20Raised%20In&vunhst_antitcellreagentcd3=Anti-T%20Cell%20Reagent%20CD3&vunhst_antitcellreagentdiluted=Anti-T%20Cell%20Reagent%20Diluted&vunhst_antitcellreagentdilutedin=Anti-T%20Cell%20Reagent%20Diluted%20In&vunhst_antitcellreagentdilution=Anti-T%20Cell%20Reagent%20Dilution&vunhst_antitcellreagentlabelledwith=Anti-T%20Cell%20Reagent%20Labelled%20with&vunhst_antitcellreagentmanufacturer=Anti-T%20Cell%20Reagent%20Manufacturer&vunhst_antitcellreagentnumberofwashes=Anti-T%20Cell%20Reagent%20Number%20of%20Washes&vunhst_antitcellreagentotherdetails=Anti-T%20Cell%20Reagent%20Other%20Details&vunhst_antitcellreagenttemperature=Anti-T%20Cell%20Reagent%20Temperature&vunhst_antitcellreagentvolumetube=Anti-T%20Cell%20Reagent%20Volume%2fTube&vunhst_antitcellreagentwashmedium=Anti-T%20Cell%20Reagent%20Wash%20Medium&vunhst_antitcellreagentwashstepdetails=Anti-T%20Cell%20Reagent%20Wash%20Step%20Details&vunhst_bcellincubationtemp=B-Cell%20Incubation%20Temp&vunhst_bcellincubationtimemins=B-Cell%20Incubation%20Time%20%28mins%29&vunhst_bcellvolumetube=B-Cell%20Volume%2fTube&vunhst_bcellsmaximum=B-Cells%20Maximum&vunhst_bcellsminimum=B-Cells%20Minimum&vunhst_cellconcentration=Cell%20Concentration&vunhst_celldiluent=Cell%20Diluent&vunhst_cellpreparationmethod=Cell%20Preparation%20Method&vunhst_cellvolumetube=Cell%20Volume%2fTube&vunhst_comments=Comments&createdby=Created%20By&createdonbehalfby=Created%20By%20%28Delegate%29&vunhst_dataacquisitioncalibrationbeadsused=Data%20Acquisition%20Calibration%20Beads%20Used&vunhst_dataacquisitionlivegated=Data%20Acquisition%20Live%20Gated&vunhst_dataacquisitionmakeofflowcytometer=Data%20Acquisition%20Make%20of%20Flow%20Cytometer&vunhst_dataacquisitionparameterassesspeakperformance=Data%20Acquisition%20Parameter%20Assess%20Peak%20Performance&vunhst_dataacquisitionpeakperformancepleasespecify=Data%20Acquisition%20Peak%20Performance%20Please%20Specify&vunhst_dataacquisitionscaleused=Data%20Acquisition%20Scale%20Used&vunhst_dataacquisitionscaleusedpleasespecify=Data%20Acquisition%20Scale%20Used%20Please%20Specify&vunhst_dataacquisitiontypeofbeads=Data%20Acquisition%20Type%20of%20Beads&vunhst_dataacquisitionusedfor=Data%20Acquisition%20Used%20For&vunhst_dataanalysisbcellminimumvalueforpcs=Data%20Analysis%20B%20Cell%20Minimum%20Value%20for%20PCS&vunhst_dataanalysisdistinguishbindingofhumanigg=Data%20Analysis%20Distinguish%20Binding%20of%20Human%20IgG&vunhst_dataanalysisotherdetails=Data%20Analysis%20Other%20Details&vunhst_dataanalysisparameterusedtoassessdifference=Data%20Analysis%20Parameter%20used%20to%20Assess%20Difference&vunhst_dataanalysispositivefromnegative=Data%20Analysis%20Positive%20From%20Negative&vunhst_dataanalysistcellminiumumvalueforpcs=Data%20Analysis%20T%20Cell%20Miniumum%20Value%20for%20PCS&vunhst_detectionstageincubationtimemins=Detection%20Stage%20%20Incubation%20Time%20%28mins%29&vunhst_detectionstagedilutedin=Detection%20Stage%20Diluted%20In&vunhst_detectionstagedilution=Detection%20Stage%20Dilution&vunhst_detectionstageincubationtemperature=Detection%20Stage%20Incubation%20Temperature&vunhst_detectionstagemoredetails=Detection%20Stage%20More%20Details&vunhst_detectionstagevolumetube=Detection%20Stage%20Volume%20%2f%20Tube&vunhst_distribution_lab=Distribution_Lab&vunhst_finalwashmedium=Final%20Wash%20Medium&vunhst_finalwashnumberofwashes=Final%20Wash%20Number%20of%20Washes&vunhst_finalwashotherdetails=Final%20Wash%20Other%20Details&vunhst_finalwashresuspensionofcells=Final%20Wash%20Resuspension%20of%20Cells&vunhst_finalwashstepdetails=Final%20Wash%20Step%20Details&vunhst_finalwashtemperaturec=Final%20Wash%20Temperature%20%28C%29&vunhst_finalwashvolumetubeml=Final%20Wash%20Volume%2fTube%20%28ml%29&vunhst_finalwashvolumetubel=Final%20Wash%20Volume%2fTube%20%28%c2%b5l%29&vunhst_firstwashstep=First%20Wash%20Step&vunhst_firstwashtemperature=First%20Wash%20Temperature&vunhst_firstwashvolumetube=First%20Wash%20Volume%2fTube&vunhst_incubationtimemins=Incubation%20Time%20%28Mins%29&vunhst_isantibcellreagentaddedwithantihumanigg=Is%20Anti%20B%20Cell%20Reagent%20added%20with%20Anti%20human%20IgG&vunhst_isantitcellreagentwithantibcellreagent=Is%20Anti%20T%20cell%20reagent%20with%20Anti-B%20cell%20reagent&vunhst_isthereawashstepbeforeaddingantibcell=Is%20there%20a%20wash%20step%20before%20adding%20Anti%20B%20Cell&vunhst_isthereawashstepbeforeaddingreagent=Is%20there%20a%20wash%20step%20before%20adding%20Reagent&vunhst_manufacturer=Manufacturer&modifiedby=Modified%20By&modifiedonbehalfby=Modified%20By%20%28Delegate%29&modifiedon=Modified%20On&vunhst_negativecontrol=Negative%20Control&vunhst_numberofcrossmatchtestreplicates=Number%20of%20Crossmatch%20Test%20Replicates&vunhst_numberofnegativereplicates=Number%20of%20Negative%20Replicates&vunhst_numberofpositivereplicates=Number%20of%20Positive%20Replicates&vunhst_numberofwashes=Number%20of%20Washes&ownerid=Owner&owningbusinessunit=Owning%20Business%20Unit&vunhst_positivecontrol=Positive%20Control&vunhst_pronasetreatment=Pronase%20Treatment&vunhst_pronasetreatmentdetails=Pronase%20Treatment%20Details&overriddencreatedon=Record%20Created%20On&vunhst_temperature=Sensitisation%20Temperature&vunhst_serumdilutions=Serum%20Dilutions&vunhst_serumvolumetube=Serum%20Volume%2fTube&statuscode=Status%20Reason&vunhst_tcellincubationtemp=T-Cell%20Incubation%20Temp&vunhst_tcellincubationtimemins=T-Cell%20Incubation%20Time%20%28mins%29&vunhst_tcellvolumetube=T-Cell%20Volume%2fTube&vunhst_tcellsmaximum=T-Cells%20Maximum&vunhst_tcellsminimum=T-Cells%20Minimum&vunhst_totaleventsmaximum=Total%20Events%20Maximum&vunhst_totaleventsminimum=Total%20Events%20Minimum&vunhst_washmedium=Wash%20Medium
Yes | No
Yes | No
Yes | No
Yes | No
Yes | No
Yes | No
Yes | No
Yes | No
Yes | No
Yes | No
Yes | No
Active | Inactive