## 2B-1-2026 T cells (G1)
2B-1-2026 G1
Distribution Date: 02/06/2026
Result summary | Serum 1 | Serum 2 | Serum 3 | Serum 4
Total distributed | 19 | 19 | 19 | 19
Total submitted | 19 | 19 | 19 | 19
Not Tested | 2 | 2 | 2 | 2
Positive | 13 | 16 | 0 | 17
Negative | 4 | 1 | 17 | 0
Equivocal | 0 | 0 | 0 | 0
% Positive | 76.47% | 94.12% | 0% | 100%
% Negative | 23.53% | 5.882% | 100% | 0%
% Equivocal | 0% | 0% | 0% | 0%
Consensus | Positive | Positive | Negative | Positive
Unacceptable | Negative, Equivocal | Negative, Equivocal | Positive, Equivocal | Negative, Equivocal
Number acceptable | 13 | 16 | 17 | 17
Number unacceptable | 4 | 1 | 0 | 0
Assessed | Yes | Yes | Yes | Yes
Laboratory Code | T-cells Serum 1 | T-cells Serum 2 | T-cells Serum 3 | T-cells Serum 4 | T-cells Serum 1 Cytometer Reading | T-cells Serum 2 Cytometer Reading | T-cells Serum 3 Cytometer Reading | T-cells Serum 4 Cytometer Reading | Units for cytometer readings | Cell Viability | Comments | T-cells Local Negative Control | T-cells Local Positive Control | T-cells Positive Result Value | Date received | Date Tested | Submitted | Distribution Year | Reason for not testing | Confirm Reason | Tested | Correct | Assessed | Tcell
Lab-1003 | Positive | Positive | Negative | Positive | 5486 | 18381 | 3108 | 23961 | Median Channel Fluorescence | 95 |  | 3184 | 4548 | 40 Linear Channel Shift | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1008 | Positive | Positive | Negative | Positive | 1.36 | 2.62 | 1.06 | 3.36 |  | 100 |  | 1498 | 12896 | >1.3 RMF | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1019 | Positive | Positive | Negative | Positive | 1.13 | 3.21 | 0.47 | 4.67 | Median Log Fluorescent Intensity | 90 |  | 0.48 | 8.18 | 1.5 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1043 | Positive | Positive | Negative | Positive | 2.06 | 5.65 | 0.93 | 8.96 |  | 999 | Cell viability not tested | 7250.53 (X-Median Mean) | 21.01 | >1.5 RMF ratio | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1055 | Positive | Positive | Negative | Positive | 169 | 258 | 103 | 320 | Median Log Channel | 100 |  | 98 | 425 | RMF >1.2 | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1079 | Positive | Positive | Negative | Positive | 1.40 | 3.04 | 0.98 | 3.82 |  | 99 | Cell viability not measured - forced to enter a value so entered 99 | 1.0 | 90.35 | 1.3 RMF | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1108 | Positive | Positive | Negative | Positive | 373 | 356 | 138 | 1021 | Log median | 99 |  | 141 | 842 | 1.5 times the mean of negative control replicate median values | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1118 | Positive | Positive | Negative | Positive | 2.0 | 4.1 | 1.0 | 5.5 | MCFN | 100 |  | 1.0 | 16.7 | RR > 1.5 | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1128 | Positive | Positive | Negative | Positive | 19700 | 35704 | 7910 | 79049 | log median x | 0 |  |  | 48502 | 1.904 * av neg | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1139 | Positive | Positive | Negative | Positive | 440 | 2801 | 133 | 2399 | Median Channel Shifts | 100 |  | Mean of 148 | 1522 | Ratio of 1.595 or greater | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1153 | Negative | Positive | Negative | Positive | 232 | 280 | 207 | 347 | median channel value | 90 |  | 208 | 540 | 40 channel shift (248) | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1174 | Positive | Positive | Negative | Positive | 838 | 1728 | 627 | 2499 | Median channel fluorescence | 75 |  | 414 | 1405 | X+3SD | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1176 | Negative | Positive | Negative | Positive | 1.5 | 8.4 | 1.0 | 6.4 | RFI | 97 |  | 0.0 Negative | 52.8 Positive | >=1.67 RFI | 2026-06-03 | 2026-06-08 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1199 | Positive | Positive | Negative | Positive | 277.7 | 649.3 | 160.7 | 932.7 | MFI | 95 |  | 160.0 | 9500.3 | RMF≥1.6 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1212 | Negative | Negative | Negative | Positive | 176 (RMF = 1.04) | 181 (RMF = 1.08) | 165 (RMF = 0.98) | 217 (RMF = 1.30) | channel values | 95 | mean of 2 replicates used | 168 | 257 (RMF = 1.53) | RMF = 1.30 X mean of Negative control | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 2 | 4 | Yes
Lab-1216 | Negative | Positive | Negative | Positive | 2428 | 3592 | 2504 | 3787 | Median x (log) | 87 |  | 2216 | 43900 | >1.5 x Trimmed mean | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1238 | Positive | Positive | Negative | Positive | 87.3 | 176 | 63 | 228 | Median FL1 | 99 |  | 45.3 | 1179 | AB serum average + 2SDs | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1239 | Not Tested | Not Tested | Not Tested | Not Tested | . | . | . | . | Median log channel | 90 |  | . | . | . | 2026-06-03 |  | Yes | 2026 | Technical failure | Technical Failure | 0 | 0 | 0 | Yes
Lab-1264 | Not Tested | Not Tested | Not Tested | Not Tested | [6 empty] | Not routinely calculated. Cells were not within the usual gating. |  |  |  | 2026-06-03 | 2026-06-04 | Yes | 2026 | Cell viability too low | Cells were not within the expected lymphocyte gates. In the clinical setting we would have requested repeat donor bloods and also performed HLA antibody testing. | 0 | 0 | 0 | Yes

## 2B-1-2026 B cells (G1)
2B-1-2026 G1
Distribution Date: 02/06/2026
Result summary | Serum 1 | Serum 2 | Serum 3 | Serum 4
Total distributed | 19 | 19 | 19 | 19
Total submitted | 19 | 19 | 19 | 19
Not Tested | 2 | 2 | 2 | 2
Positive | 17 | 16 | 1 | 17
Negative | 0 | 1 | 16 | 0
Equivocal | 0 | 0 | 0 | 0
% Positive | 100% | 94.12% | 5.882% | 100%
% Negative | 0% | 5.882% | 94.12% | 0%
% Equivocal | 0% | 0% | 0% | 0%
Consensus | Positive | Positive | Negative | Positive
Unacceptable | Negative, Equivocal | Negative, Equivocal | Positive, Equivocal | Negative, Equivocal
Number acceptable | 17 | 16 | 16 | 17
Number unacceptable | 0 | 1 | 1 | 0
Assessed | Yes | Yes | Yes | Yes
Laboratory Code | B-cells Serum 1 | B-cells Serum 2 | B-cells Serum 3 | B-cells Serum 4 | B-cells Serum 1 Cytometer Reading | B-cells Serum 2 Cytometer Reading | B-cells Serum 3 Cytometer Reading | B-cells Serum 4 Cytometer Reading | Units for cytometer readings | Cell Viability | Comments | B-cells Local Negative Control | B-cells Local Positive Control | B-cells Positive Result Value | Date received | Date Tested | Submitted | Distribution Year | Reason for not testing | Confirm Reason | Tested | Correct | Assessed | Bcell
Lab-1003 | Positive | Positive | Negative | Positive | 121754 | 72567 | 11196 | 141952 | Median Channel Fluorescence | 95 |  | 19228 | 103362 | 40 Linear Channel Shift | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1008 | Positive | Positive | Negative | Positive | 5.69 | 2.76 | 1.00 | 5.42 |  | 100 |  | 2754 | 26142 | >1.3 RMF | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1019 | Positive | Positive | Negative | Positive | 19.50 | 12.44 | 9.92 | 19.45 | Median Log Fluorescent Intensity | 90 |  | 3.37 | 11.62 | 3 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1043 | Positive | Positive | Negative | Positive | 3.58 | 2.79 | 1.31 | 4.69 |  | 999 | Cell viability not tested | 53035.37 (X-Median Mean) | 12.85 | >2.0 RMF ratio | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1055 | Positive | Positive | Negative | Positive | 1857 | 796 | 219 | 1570 | Median Log Channel | 100 |  | 301 | 3973 | RMF >1.3 | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1079 | Positive | Positive | Negative | Positive | 2.71 | 2.12 | 1.21 | 3.86 |  | 99 | Cell viability not measured - forced to enter a value so entered 99 | 1.0 | 99.66 | 1.5 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1108 | Positive | Positive | Negative | Positive | 2822 | 925 | 331 | 3003 | Log median | 99 |  | 347 | 6363 | 1.5 times the mean of negative control replicate median values | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1118 | Positive | Positive | Negative | Positive | 7.8 | 4.4 | 1.0 | 9.4 | MCFN | 100 |  | 1.0 | 17.3 | RR > 2.0 | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1128 | Positive | Positive | Negative | Positive | 224350 | 161354 | 81560 | 300487 | log median x | 0 |  |  | 725424 | 2.795 * av neg control | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1139 | Positive | Positive | Negative | Positive | 3492 | 2777 | 437 | 4335 | Median Channel Shifts | 100 |  | mean of 453 | 12244 | Ratio of 2.545 or greater | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1153 | Positive | Positive | Negative | Positive | 507 | 440 | 381 | 516 | median channel value | 90 |  | 400 | 703 | 40 channel shift (440) | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1174 | Positive | Positive | Negative | Positive | 2931 | 2772 | 718 | 7857 | Median channel fluorescence | 75 |  | 604 | 9490 | X+3SD | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1176 | Positive | Positive | Negative | Positive | 1.8 | 2.9 | 0.9 | 3.2 | RFI | 97 |  | 0.0 Negative | 11.4 Positive | >=1.61 | 2026-06-03 | 2026-06-08 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1199 | Positive | Positive | Negative | Positive | 1792.3 | 916.0 | 247.7 | 2239.7 | MFI | 95 |  | 283.3 | 11394.7 | RMF≥1.6 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1212 | Positive | Negative | Negative | Positive | 318 (RMF = 1.94) | 180 (RMF = 1.10) | 163 (RMF = 0.99) | 282 (RMF = 1.72) | channel values | 95 | mean of 2 replicates used | 164 | 698 (RMF = 4.27) | RMF = 1.30 x Mean of Negative control | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1216 | Positive | Positive | Negative | Positive | 11392 | 9507 | 4367 | 10855 | Median x (log) | 87 |  | 4380 | 101721 | >3.0 x Trimmed mean | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1238 | Positive | Positive | Positive | Positive | 482.3 | 456.3 | 141 | 779.5 | Median FL1 | 99 |  | 104.5 | 3650 | AB serum average + 2SDs | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1239 | Not Tested | Not Tested | Not Tested | Not Tested | . | . | . | . | Median log channel | 90 |  | . | . | . | 2026-06-03 |  | Yes | 2026 | Technical failure | Technical Failure | 0 | 0 | 0 | Yes
Lab-1264 | Not Tested | Not Tested | Not Tested | Not Tested | [6 empty] | Not routinely calculated. Cells were not within the usual gating. |  |  |  | 2026-06-03 | 2026-06-04 | Yes | 2026 | Cell viability too low | Cells were not within the expected lymphocyte gates. In the clinical setting we would have requested repeat donor bloods and also performed HLA antibody testing. | 0 | 0 | 0 | Yes

## 2B-2-2026 T cells (G1)
2B-2-2026 G1
Distribution Date: 02/06/2026
Result summary | Serum 1 | Serum 2 | Serum 3 | Serum 4
Total distributed | 19 | 19 | 19 | 19
Total submitted | 19 | 19 | 19 | 19
Not Tested | 2 | 2 | 2 | 2
Positive | 13 | 0 | 16 | 1
Negative | 4 | 17 | 1 | 16
Equivocal | 0 | 0 | 0 | 0
% Positive | 76.47% | 0% | 94.12% | 5.882%
% Negative | 23.53% | 100% | 5.882% | 94.12%
% Equivocal | 0% | 0% | 0% | 0%
Consensus | Positive | Negative | Positive | Negative
Unacceptable | Negative, Equivocal | Positive, Equivocal | Negative, Equivocal | Positive, Equivocal
Number acceptable | 13 | 17 | 16 | 16
Number unacceptable | 4 | 0 | 1 | 1
Assessed | Yes | Yes | Yes | Yes
Laboratory Code | T-cells Serum 1 | T-cells Serum 2 | T-cells Serum 3 | T-cells Serum 4 | T-cells Serum 1 Cytometer Reading | T-cells Serum 2 Cytometer Reading | T-cells Serum 3 Cytometer Reading | T-cells Serum 4 Cytometer Reading | Units for cytometer readings | Cell Viability | Comments | T-cells Local Negative Control | T-cells Local Positive Control | T-cells Positive Result Value | Date received | Date Tested | Submitted | Distribution Year | Reason for not testing | Confirm Reason | Tested | Correct | Assessed | Tcell
Lab-1003 | Positive | Negative | Positive | Negative | 6396 | 3455 | 27985 | 4248 | Median Channel Fluorescence | 95 |  | 2998 | 7392 | 40 Linear Channel Shift | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1008 | Positive | Negative | Positive | Negative | 1.44 | 0.85 | 5.05 | 0.91 |  | 100 |  | 2468 | 56318 | >1.3 RMF | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1019 | Positive | Negative | Positive | Negative | 0.92 | 0.49 | 5.39 | 0.54 | Median Log Fluorescent Intensity | 90 |  | 0.46 | 13.14 | 1.5 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1043 | Positive | Negative | Positive | Negative | 1.58 | 0.95 | 6.84 | 1.19 |  | 999 | Cell viability not tested | 6525.3 (X Median Mean) | 68.94 | >1.5 RMF ratio | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1055 | Positive | Negative | Positive | Positive | 130 | 104 | 507 | 118 | Median Log Channel | 100 |  | 94 | 1254 | RMF >1.2 | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1079 | Negative | Negative | Positive | Negative | 1.15 | 0.84 | 5.46 | 0.81 |  | 99 | Cell viability not measured - forced to enter reading | 1.0 | 101.84 | 1.3 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1108 | Positive | Negative | Positive | Negative | 265 | 190 | 703 | 212 | Log Median | 99 |  | 186 | 738 | 1.5 times the mean of negative control replicate median values | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1118 | Positive | Negative | Positive | Negative | 1.7 | 0.9 | 7.8 | 1.0 | MCFN | 95 | B-cells not tested - unable to achieve sufficient cell yield from the EDTA provided - T-cells acceptable. | 1.0 | 9.6 | RR > 1.5 | 2026-06-03 | 2026-06-04 | Yes | 2026 | Cell count too low | B-cells unable to be accepted due to very low cell yield. T-cells acceptable. | 4 | 4 | 4 | Yes
Lab-1128 | Negative | Negative | Positive | Negative | 11967 | 7672 | 44807 | 9312 | log median x | 0 |  |  | 45929 | 1.904 * av neg | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1139 | Positive | Negative | Positive | Negative | 333 | 165 | 598 | 221 | Median Channel Shift | 100 | Sample 2B 02 2026 acquired on the FACSCanto flow cytometer | Mean of 191 | 567 | A MCS difference of 50 or more | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1153 | Positive | Negative | Positive | Negative | 244 | 201 | 323 | 203 | median channel value | 90 |  | 204 | 623 | 40 channel shift (244) | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1174 | Positive | Negative | Positive | Negative | 453 | 239 | 1790 | 251 | Median channel fluorescence | 75 |  | 208 | 1765 | X+3SD | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1176 | Positive | Negative | Positive | Negative | 2.0 | 0.9 | 19.5 | 1.2 | RFI | 95 |  | 0.0 Negative | 30.9 Positive | >=1.67 | 2026-06-03 | 2026-06-08 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1199 | Positive | Negative | Positive | Negative | 271.7 | 149.7 | 867.7 | 161.7 | MFI | 95 |  | 155.7 | 9236.0 | RMF≥1.6 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1212 | Negative | Negative | Negative | Negative | 193 (RMF = 1.03) | 185 (RMF = 0.99) | 219 (RMF = 1.17) | 182 (RMF = 0.97) | channel values | 95 |  | 187 | 507 (RMF = 2.72) | RMF = 1.30 X mean of Negative control | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 2 | 4 | Yes
Lab-1216 | Negative | Negative | Positive | Negative | 1771 | 1734 | 2523 | 1775 | Median x (log) | 87 |  | 1567 | 13851 | >1.5 x Trimmed mean | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1238 | Positive | Negative | Positive | Negative | 122.7 | 83 | 351.7 | 88.7 | Median FL1 | 99 |  | 94 | 2519 | AB serum average + 2SDs | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1239 | Not Tested | Not Tested | Not Tested | Not Tested | . | . | . | . | Median log channel | 90 |  | . | . | . | 2026-06-03 |  | Yes | 2026 | Technical failure | Technical failure | 0 | 0 | 0 | Yes
Lab-1264 | Not Tested | Not Tested | Not Tested | Not Tested | [6 empty] | cells did not fall within the expected lymphocyte gates. |  |  |  | 2026-06-03 | 2026-06-04 | Yes | 2026 | Cell viability too low | Cells did not fall within the expected lymphocyte gates. In the clinical setting we would request repeat donor bloods and perform HLA antibody testing. | 0 | 0 | 0 | Yes

## 2B-2-2026 B cells (G1)
2B-2-2026 G1
Distribution Date: 02/06/2026
Result summary | Serum 1 | Serum 2 | Serum 3 | Serum 4
Total distributed | 19 | 19 | 19 | 19
Total submitted | 19 | 19 | 19 | 19
Not Tested | 3 | 3 | 4 | 3
Positive | 16 | 16 | 13 | 0
Negative | 0 | 0 | 2 | 16
Equivocal | 0 | 0 | 0 | 0
% Positive | 100% | 100% | 86.67% | 0%
% Negative | 0% | 0% | 13.33% | 100%
% Equivocal | 0% | 0% | 0% | 0%
Consensus | Positive | Positive | Positive | Negative
Unacceptable | Negative, Equivocal | Negative, Equivocal | Negative, Equivocal | Positive, Equivocal
Number acceptable | 16 | 16 | 13 | 16
Number unacceptable | 0 | 0 | 2 | 0
Assessed | Yes | Yes | Yes | Yes
Laboratory Code | B-cells Serum 1 | B-cells Serum 2 | B-cells Serum 3 | B-cells Serum 4 | B-cells Serum 1 Cytometer Reading | B-cells Serum 2 Cytometer Reading | B-cells Serum 3 Cytometer Reading | B-cells Serum 4 Cytometer Reading | Units for cytometer readings | Cell Viability | Comments | B-cells Local Negative Control | B-cells Local Positive Control | B-cells Positive Result Value | Date received | Date Tested | Submitted | Distribution Year | Reason for not testing | Confirm Reason | Tested | Correct | Assessed | Bcell
Lab-1003 | Positive | Positive | Positive | Negative | 153068 | 58208 | 57040 | 14292 | Median Channel Fluorescence | 95 |  | 17045 | 71491 | 40 Linear Channel Shift | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1008 | Positive | Positive | Positive | Negative | 7.67 | 3.52 | 2.91 | 0.92 |  | 100 |  | 5528 | 105089 | >1.3 RMF | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1019 | Positive | Positive | Positive | Negative | 20.77 | 16.52 | 11.39 | 8.4 | Median Log Fluorescent Intensity | 90 |  | 3.62 | 47.62 | 3 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1043 | Positive | Positive | Positive | Negative | 5.03 | 2.87 | 2.29 | 1.91 |  | 999 | Cell viability not tested | 38970.43 (X Median Mean) | 26.29 | >2.0 RMF ratio | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1055 | Positive | Positive | Positive | Negative | 3005 | 1151 | 1138 | 306 | Median Log Channel | 100 |  | 445 | 3445 | RMF >1.3 | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1079 | Positive | Positive | Positive | Negative | 4.57 | 2.84 | 2.21 | 1.07 |  | 99 | Cell viability not measured - forced to enter reading | 1.0 | 154.24 | 1.5 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1108 | Positive | Positive | Positive | Negative | 3965 | 1775 | 1385 | 516 | Log Median | 99 |  | 434 | 6162 | 1.5 times the mean of negative control replicate median values | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1118 | Not Tested | Not Tested | Not Tested | Not Tested | N/A | N/A | N/A | N/A | MCFN | 95 | B-cells not tested - unable to achieve sufficient cell yield from the EDTA provided - T-cells acceptable. | N/A | N/A | RR > 2.0 | 2026-06-03 | 2026-06-04 | Yes | 2026 | Cell count too low | B-cells unable to be accepted due to very low cell yield. T-cells acceptable. | 0 | 0 | 0 | Yes
Lab-1128 | Positive | Positive | Positive | Negative | 280292 | 147761 | 117862 | 63642 | log median x | 0 |  |  | 419771 | 2.795 * av neg control | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1139 | Positive | Positive | Positive | Negative | 547 | 444 | 475 | 272 | Median Channel Shift | 100 | Sample 2B 02 2026 acquired on the FACSCanto flow cytometer | Mean of 272 | 593 | a Median Channel shift of 100 or greater | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1153 | Positive | Positive | Negative | Negative | 544 | 511 | 466 | 436 | median channel value | 90 |  | 450 | 763 | 40 channel shift (490) | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1174 | Positive | Positive | Positive | Negative | 6721 | 1721 | 2777 | 655 | Median channel fluorescence | 75 |  | 452 | 4018 | X+3SD | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1176 | Positive | Positive | Positive | Negative | 3.5 | 1.6 | 2.5 | 1.0 | RFI | 95 |  | 0.0 Negative | 7.5 Positive | >=1.61 | 2026-06-03 | 2026-06-08 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1199 | Positive | Positive | Positive | Negative | 2283.7 | 1161.0 | 779.3 | 277.3 | MFI | 95 |  | 396.7 | 10629.0 | RMF≥1.6 | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1212 | Positive | Positive | Negative | Negative | 358 (RMF = 2.05) | 255 (RMF = 1.46) | 185 (RMF = 1.06) | 162 (RMF = 0.93) | channel values | 95 |  | 175 | 1088 (RMF = 6.23) | RMF = 1.30 x Mean of Negative control | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 4 | 3 | 4 | Yes
Lab-1216 | Positive | Positive | Not Tested | Negative | 16495 | 9146 |  | 3054 | Median x (log) | 87 |  | 3136 | 48769 | >3.0 x Trimmed mean | 2026-06-03 | 2026-06-03 | Yes | 2026 |  |  | 3 | 3 | 3 | Yes
Lab-1238 | Positive | Positive | Positive | Negative | 856.3 | 449.7 | 471 | 251.3 | Median FL1 | 99 |  | 256.5 | 4661 | AB serum average + 2SDs | 2026-06-03 | 2026-06-04 | Yes | 2026 |  |  | 4 | 4 | 4 | Yes
Lab-1239 | Not Tested | Not Tested | Not Tested | Not Tested | . | . | . | . | Median log channel | 90 |  | . | . | . | 2026-06-03 |  | Yes | 2026 | Technical failure | Technical failure | 0 | 0 | 0 | Yes
Lab-1264 | Not Tested | Not Tested | Not Tested | Not Tested | [6 empty] | cells did not fall within the expected lymphocyte gates. |  |  |  | 2026-06-03 | 2026-06-04 | Yes | 2026 | Cell viability too low | Cells did not fall within the expected lymphocyte gates. In the clinical setting we would request repeat donor bloods and perform HLA antibody testing. | 0 | 0 | 0 | Yes

## hiddenSheet (hidden)
vunhst_results_scheme2b:4JcOrnhthYxmR+lZWqzn2Ze2ZggyMxUPVvGxcWbnLfm/5Zw2FDUcrwIIWB4UTScfi9T8RctGp9lq41azsXa1ng==:vunhst_results_scheme2bid=%28Do%20Not%20Modify%29%20Results_Scheme2B&checksumLogicalName=%28Do%20Not%20Modify%29%20Row%20Checksum&modifiedon=%28Do%20Not%20Modify%29%20Modified%20On&vunhst_sample=Sample&vunhst_laboratory=Laboratory&4ba235cb-fc6c-4ea9-b824-af5c417c0441.vunhst_laboratorycode=Laboratory%20Code%20%28Laboratory%29%20%28Laboratory%29&vunhst_datereceived=Date%20received&vunhst_datetested=Date%20Tested&vunhst_issubmitted=isSubmitted&ae01c472-d563-446a-a24d-03d56717b73f.vunhst_distributionyear=Distribution%20Year%20%28Distribution%20Lab%29%20%28Distribution_Lab%29&vunhst_bcellslocalnegativecontrol=B-cells%20Local%20Negative%20Control&vunhst_bcellslocalpositivecontrol=B-cells%20Local%20Positive%20Control&vunhst_bcellspositiveresultvalue=B-cells%20Positive%20Result%20Value&vunhst_bcellsserum1cytometerreading=B-cells%20Serum%201%20Cytometer%20Reading&vunhst_bcellsserum1=B-cells%20Serum%201&vunhst_bcellsserum2=B-cells%20Serum%202&vunhst_bcellsserum2cytometerreading=B-cells%20Serum%202%20Cytometer%20Reading&vunhst_bcellsserum3=B-cells%20Serum%203&vunhst_bcellsserum3cytometerreading=B-cells%20Serum%203%20Cytometer%20Reading&vunhst_bcellsserum4cytometerreading=B-cells%20Serum%204%20Cytometer%20Reading&vunhst_bcellsserum4=B-cells%20Serum%204&vunhst_cellviability=Cell%20Viability&vunhst_comments=Comments&vunhst_confirmreason=Confirm%20Reason&vunhst_reasonfornottestingnew=Reason%20for%20not%20testing&vunhst_tcellslocalnegativecontrol=T-cells%20Local%20Negative%20Control&vunhst_tcellslocalpositivecontrol=T-cells%20Local%20Positive%20Control&vunhst_tcellspositiveresultvalue=T-cells%20Positive%20Result%20Value&vunhst_tcellsserum1=T-cells%20Serum%201&vunhst_tcellsserum1cytometerreading=T-cells%20Serum%201%20Cytometer%20Reading&vunhst_tcellsserum2=T-cells%20Serum%202&vunhst_tcellsserum2cytometerreading=T-cells%20Serum%202%20Cytometer%20Reading&vunhst_tcellsserum3=T-cells%20Serum%203&vunhst_tcellsserum3cytometerreading=T-cells%20Serum%203%20Cytometer%20Reading&vunhst_tcellsserum4=T-cells%20Serum%204&vunhst_tcellsserum4cytometerreading=T-cells%20Serum%204%20Cytometer%20Reading&vunhst_unitsforcytometerreadings=Units%20for%20cytometer%20readings
Yes | No