## 2B-1-2-2026 G2 Methods
| Cell Preparation and Controls | [12 empty] | Sensitisation Stage |  | First Wash Step |  |  |  |  | Detection Stage | [8 empty] | Anti-T Cell Reagent | [17 empty] | Anti-B-cell Reagents | [17 empty] | Final Wash | [7 empty] | Data Acquisiton | [14 empty] | Data Analysis
Laboratory ID | Cell Preparation Method | Cell Volume/Tube | Cell Concentration | Cell Diluent | Serum Volume/Tube | Pronase Treatment | Pronase Treatment Details | Serum Dilutions | Negative Control | Number of Negative Replicates | Positive Control | Number of Positive Replicates | Number of Crossmatch Test Replicates | Incubation Time (Mins) | Sensitisation Temperature | First Wash Step | Wash Medium | First Wash Volume/Tube | Number of Washes | First Wash Temperature | Manufacturer | Anti-human IgG Reagent Raised In | Anti-human IgG Labelled With | Detection Stage More Details | Detection Stage Dilution | Detection Stage Diluted In | Detection Stage Volume / Tube | Detection Stage Incubation Time (mins) | Detection Stage Incubation Temperature | Anti-T Cell Reagent Manufacturer | Anti-Human IgG Reagent added with Anti-T cell | Is there a wash step before adding Reagent | Anti-T Cell Reagent Wash Step Details | Anti-T Cell Reagent Wash Medium | T-Cell Volume/Tube | Anti-T Cell Reagent Number of Washes | T-Cell Incubation Temp | Is Anti T cell reagent with Anti-B cell reagent | Anti-T Cell Reagent Labelled with | Anti-T Cell Reagent Other Details | Anti-T Cell Reagent CD3 | Anti-T Cell Reagent Diluted | Anti-T Cell Reagent Diluted In | Anti-T Cell Reagent Dilution | Anti-T Cell Reagent Volume/Tube | T-Cell Incubation Time (mins) | Anti-T Cell Reagent Temperature | Anti B-Cell Reagent Manufacturer | Is Anti B Cell Reagent added with Anti human IgG | Is there a wash step before adding Anti B Cell | Anti B-Cell Reagent Wash Step Details | Anti B-Cell Reagent Wash Medium | B-Cell Volume/Tube | Anti B-Cell Reagent Number of Washes | Anti B-Cell Reagent Temperature | Anti B-Cell Reagent Labelled With | Anti B-Cell Reagent Other Details | Anti B-Cell Reagent CD19 | Anti B-Cell Reagent CD20 | Anti B-Cell Reagent Diluted | Anti B-Cell Reagent Diluted In | Anti B-Cell Reagent Dilution | Anti B-Cell Reagent Volume/Tube | B-Cell Incubation Time (mins) | B-Cell Incubation Temp | Final Wash Step Details | Final Wash Medium | Final Wash Volume/Tube (ml) | Final Wash Number of Washes | Final Wash Temperature (C) | Final Wash Resuspension of Cells | Final Wash Other Details | Final Wash Volume/Tube (µl) | Data Acquisition Make of Flow Cytometer | Data Acquisition Live Gated | Data Acquisition Calibration Beads Used | Data Acquisition Used For | Data Acquisition Type of Beads | Total Events Minimum | Total Events Maximum | T-Cells Minimum | T-Cells Maximum | B-Cells Minimum | B-Cells Maximum | Data Acquisition Parameter Assess Peak Performance | Data Acquisition Peak Performance Please Specify | Data Acquisition Scale Used | Data Acquisition Scale Used Please Specify | Data Analysis Parameter used to Assess Difference | Data Analysis Other Details | Data Analysis T Cell Miniumum Value for PCS | Data Analysis B Cell Minimum Value for PCS | Data Analysis Positive From Negative | Data Analysis Distinguish Binding of Human IgG | Comments
Lab-1013 | Gradient Density Separation | 30 | 16x10^6 | PBA | 50 | No |  | No | Merck Human AB Serum | 3 | Campath | 2 | 3 | 30 | 22 | Manual | PBA | 2500 | 2 | RT | Dako | Rabbit | FITC |  | 1:20 | PBA | 5 | 30 | 4 | BD | Yes |  | Manual | PBA | 5 | 2 | 4 | Yes | Other | Peridinin Chlorophyll Protein (PerCP) | Yes | No |  |  | 2500 | 30 | RT | BD | Yes |  | Manual | PBA | 5 | 2 | 4 | Phycoerythrin |  | Yes | No | No |  |  | 2500 | 30 | 4 | Manual | PBA | 2500 | 2 | RT | Sheath fluid |  | 250 | BD FACS LYRIC | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check | CS&T | 10000 | NA | 800 | NA | 800 | NA | Median |  | Other | Biexponential | Median difference |  | ≥65 Channel Shift | ≥125 Channel Shift
Lab-1014 | density gradient | 35 | 5 000 000 cells/ml | PBS 1X | 15 | No |  | No | NIBSC | 3 | HUMAN CD52 BIOTINYLATED ANTIBODY (réf FAB9889B)- R et D SYSTEMS | 3 | 2 | 20 | room temperature | Manual | PBS/BSA à 2% | 150 or 200 | 2 | room temperature | Jackson Immuno Research | Goat | Fab2 fragment |  | 1/10 | water quality molecular biology | 1.25 | 10 | room temperature | BD | Yes |  | Manual | PBS/BSA à 2% | 5 | 2 | room temperature | Yes |  | APC | Yes | No |  |  | 150 or 200 | 10 | room temparature | BD | Yes |  | Manual | PBS/BSA à 2% | 1.25 | 2 | room tempature |  | Phycoerythrin-cyanine 7 | Yes | No | No |  |  | 150 or 200 | 10 | room temperature | Manual |  |  |  |  | Sheath fluid |  | 200 | BD FACSLyric | No | Yes | [8 empty] | Median | [5 empty] | 1,4 (ratio) | 1,6 (ratio) | Ly T 1,3 (ratio) ; LyB 1,4 (ratio)
Lab-1024 | Lymphocytes isolation | 200 | 250000 | PBS+4% goat serum | 20 | Yes | 20mg/ml | yes 1/2 | 3 healthy male donnors without transfusions | 3 | innotrain | 1 | 1 | 30 | 4 | Manual | PBS+4% goat serum | 0.2 | 3 | 4 | BD + Dako schweiz | Goat | FITC |  | none | none | 0.9 | 20 | 4 | BD | Yes |  | Manual | PBS+4% goat serum | 0.5 | 3 | 4 | Yes | Other | APC | Yes | No |  |  | 0.2 | 20 | 4 | BD | Yes |  | Manual | PBS+4% goat serum | 2.5 | 3 | 4 | Phycoerythrin |  | Yes | No | No |  |  | 0.2 | 20 | 4 | Manual | PBS | 0.2 | 1 | 4 | Phosphate Buffered Saline |  | 500 | BD FACSLyric | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | BD Cytometer Setup and Tracking Beads | [6 empty] | Median |  | Other | biexpnential | Median difference |  |  |  |  | Mean of 3 negative contols median + 3 SD
Lab-1032 | Density gradient centrifugation | 15μl | 10e+6/ml | PBS with FBS (2%) | 30μl/tube | Yes | 4.7 U/ml | no | Pooled negative serum (patients) | 2 | Pooled positive serum (patients) | 1 | 2 | 20 minutes | 24°C | Manual | PBS with FBS (2%) | 0.15 ml | 4 | 24°C | Jackson Laboratory | Goat | FITC |  | 1/1000 | PBS with FBS (2%) | 0.05 μl/tube | 5 min | 24°C | BD | Yes | No | Manual |  | 2 μl |  | 24°C | Yes | Other | PerCP | Yes | Yes | PBS with FBS (2%) | 2 μl in 50 μl |  | 5 min |  | BD | Yes | No | Manual |  | 1 μl/tube |  |  | Other | APC | Yes | No | Yes | PBS with FBS (2%) | 1 μl in 50 μl |  | 5 | 24°C | Manual | PBS with FBS (2%) | 0.15 ml | 1 | 24°C | Other | PBS with FBS (2%) | 200 μl/tube | FACSLyric | No | Yes | Alignment Check, Laser Check, Fluorescence Check | BD™ CS&T Beads |  |  |  |  | 600 | 600 | Median |  | Log |  | Median difference |  | 280 channel shift | 400 channel shift
Lab-1035 | Ficoll gradient | 50 | 5x10E5/well | RPMI + 1% BSA | 25 | No |  | No | Phlebectomy | 2 | ETRL | 2 | 2 | 30 | 20 | Manual | PBS | 0,1 | 3 | 20 | Jackson Immuno Research | Goat | FITC |  | 400x | PBS | 25 | 15 | 20 | Becton Dickinson | Yes |  | Manual | PBS | 25 | 2 | 20 | Yes | Other | BV786 | Yes | Yes | PBS | 50x | 0,1 | 15 | 20 | Becton Dickinson | Yes |  | Manual | PBS | 25 | 2 | 20 | Other | APC | Yes | No | Yes | PBS | 50x | 0,1 | 15 | 20 | Manual | PBS | 0,1 | 2 (same wash as after the anti-T and anti-B detection) | 20 | Phosphate Buffered Saline |  | 150 | BD FACSLyric | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | CS&T beads |  | 10000 | 1000 |  | 500 |  | Median |  | Log |  | Other | MFI ratio compared to negative control | Low positive control ratio >1,3 and high positive control ratio > 10 | Low positive control ratio >2,8 and high positive control ratio > 10 | Ratio >1,3 for T cells and ratio > 2,8 for B cells | Ratio >1,3 for T cells and ratio > 2,8 for B cells
Lab-1040 | [5 empty] | No | [24 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No
Lab-1047 | EasySep method | 15 µl | 1.10^6 | PBS | 30 | Yes | 1 mg/ml | No | Known negative serum | / | / | / | / | 20 | Ambiant | Manual | RPMI | 0,2 | 3 | Ambiant | FACS Lyric | Goat | FITC | / |  | Distilled water | 10 | 5 | Ambiant |  | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [5 empty] | Manual | PBS | 0,2 | 2 | Ambiant | Phosphate Buffered Saline |  | 400 | Kaluza | Yes | Yes | Alignment Check, Laser Check | [7 empty] | Other | Median of linear value | Linear |  |  |  | 130 | 150
Lab-1054 | EasySep Kit from Stemcell Technologies | 15 | 10 x 10^6 cells/mL | DPBS | 30 | Yes | 2 mg/mL | yes | in house | 1 | in house | 1 | 2 | 20 | 23 | Manual | DPBS + 2 % FCS | 0.2 | 3 | 23 | Jackson Immuno Research | Goat | FITC |  | 1.5 mg/mL | H2O | 0.5 | 10 | 23 | BD | Yes |  | Manual | DPBS + 2 % FCS | 2 | 2 | 23 | Yes | Other | PerCP | Yes | No |  |  | 200 | 10 | 23 | biolegend | Yes |  | Manual | DPBS + 2 % FCS | 1.5 | 2 | 23 | Other | APC-Cy7 | Yes | No | No |  |  | 200 | 10 | 23 | Manual | DPBS + 2 % FCS | 200 | 2 | 23 | Other | DPBS + 2 % FCS | 220 | BD FACSLyric | Yes | Yes | Laser Check, Fluorescence Check | CS&T beads | 10 000 | 20 000 |  |  |  |  | Median |  | Linear, Log |  | Other | Ratio: patient/negative control | 3.6 | 3.8
Lab-1060 | [5 empty] | No | [24 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No
Lab-1061 | [5 empty] | Yes | [24 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No | [18 empty] | No methods provided. Not tested due to sample delivery delay.
Lab-1067 | Beadseparated cells | 10 | 10^6 cells/mL | PBS/25 FBS | 20 | Yes | 0.2 mg/mL | No | Non immunized male blood donors | 2 | anti CD52 | 1 | 2 | 30 | 20 | Manual | PBS/2% FBS | 0.150 | 2 | 4 | Jackson | Goat | FITC |  | 1:100 | PBS/2% FBS | 20 | 10 | 4 | BD Biossciences | Yes |  | Manual | PBS/2% FBS |  | 2 |  | No | Phycoerythrin |  | Yes | Yes |  |  | 150 |  | 4 | BD Biosciences | Yes |  | Manual |  |  |  |  | Phycoerythrin |  | Yes | No | No | [5 empty] | Manual | PBS/2% FBS | 150 | 2 | 4 | Other | PBS/2% FBS | 200 | DxFlex (Beckman Coulter) | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | DxFlex Daily QC Fluorophores, Ref C39283 | NA | NA | 500 | NA | 500 | NA | Median |  | Log |  | Other | Ratio | 20 | 20 | 3.4
Lab-1070 | [5 empty] | No | [24 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No
Lab-1071 | [5 empty] | No | [24 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No
Lab-1075 | EasySep | 100 | 3,0millj/ml | 2%FCS in 1xPBS | 50 | No |  | No | NIBSC | 2 | Innotrain | 2 | 2 | 30 | 22 | Manual | 2%FCS in 1xPBS | 2 | 3 | 22 | Jackson ImmunoResearch | Goat | FITC |  | 1:4 | 2%FCS in 1xPBS | 10 | 20 | 4 | Becton Dickinson | Yes |  | Manual | 2%FCS in 1xPBS | 10 | 2 | 4 | Yes | Other | PerCP | No | Yes | 2%FCS in 1xPBS | 1:2 | 2 | 20 | 4 | Becton Dickinson | Yes |  | Manual | 2%FCS in 1xPBS | 10 | 2 | 4 | Phycoerythrin |  | Yes | No | Yes | 2%FCS in 1xPBS | 1:2 | 2 | 20 | 4 | Manual | 2%FCS in 1xPBS | 2 | 2 | 4 | Other | 2%FCS in 1xPBS | 350 | Becton Dickinson FACS Lyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | BD Trucount Tubes |  |  | 500 |  | 500 |  | Median |  | Log |  | Median difference |  | 15.000MFI | 15.000MFI | 1,5 x Neg control
Lab-1076 |  | 20 | 300 x10^3 cells in 20 ul | PBS | 40 | No |  | n | POOLED (negative PRA; >20 sera) | 2 | POOLED (>20 Positive diluted in negative sera) | 2 | 2 | 20 | room | Manual | pbs | 150 ul | 3 | room | jackson lab | Goat | Phycoerythrin |  | y | PBS | 3/50 | 20 | ROOM |  | Yes |  | Manual | PBS |  | 3 |  | Yes | Other | PerCP | Yes | No |  |  | 150 ul/well |  | room |  | Yes | [6 empty] | FITC |  | Yes | No | No | [5 empty] | Manual |  |  |  |  | Phosphate Buffered Saline |  | 300ul/tube | BD LYRIC | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check | [5 empty] | 3000 |  | Mean |  | Log |  | Mean difference
Lab-1092 | Ficoll | 20 | 2x10^5 | 1% BSA in PBS | 20 | No |  | No | In house | 2 | ETRL | 1 | 2 | 30 | RT | Manual | 1% BSA in PBS | 0.2 | 3 | 4 | DAKO/Agilent | Rabbit | FITC |  | 1:50 | PBS | 20 | 30 | 4 | BD Biosciences | Yes |  | Manual | 1% BSA in PBS | 20 | 2 | 4 | Yes | Other | V500 | Yes | Yes | PBS | 1:25 | 0.2 | 30 | 4 | DAKO/Agilent | Yes |  | Manual | 1% BSA in PBS | 20 | 2 | 4 | Other | APC | Yes | No | Yes | PBS | 1:25 | 0.2 | 30 | 4 | Manual | 1% BSA in PBS | 0.2 | 2 | 4 | Other | 1% BSA in PBS | 150 | BD Lyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | BD Cytometer Setup & Tracking Beads | 20000 | [5 empty] | Median |  | Log |  | Other | ratio of median fluorescence sample / median fluorescence neg ctrl | Ratio >20 | Ratio >20 | T cells: ratio >1.17; B cells: ratio >1.21 | n.a.
Lab-1095 | NA |  |  |  |  | No | [7 empty] | NA | [15 empty] | NA | No | [6 empty] | No |  |  | No | No | [5 empty] | NA | No | [8 empty] | No | No | No | [14 empty] | No | No
Lab-1096 | EASYSEP DIRECT HUM LYMPHOCYTE ISOLATION | 100 | 3500 cell/ul | serum physiologique | 20 | Yes | 2,35 U/mL | no | Control qualité interne | 2 | Control qualité interne | 2 | 2 | 25 | RT | Manual | PBS | 150 | 2 | RT | DxFlex Beckman Coulter | Goat | FITC |  | 1/50 | PBS | 50 | 15 | 4°C | Beckman | Yes | No | Manual |  | 50 |  | 4 | Yes |  | APC-A750 | Yes | Yes | PBS | 1/10 |  | 15 |  | Beckman | Yes | No | Manual |  | 50 |  |  | Other | SNv428 | Yes | No | Yes | PBS | 1/10 |  | 15 | 4 | Manual | PBS | 120 | 1 | RT | Phosphate Buffered Saline |  | 110 | DxFlex | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | DXFLEX DAILY QC FLUOROSPHERES | [6 empty] | Median |  | Log |  |  | ratio Median Sample / Median Control | 1,6 | 1,8
Lab-1111 | Immunomagnetic negative isolation directly from whole blood, EasySep Direct Human Total Lymphocyte Isolation Kit—Stemcell Technologies, Catalog # 19655 | 10 ul per well | 10.0 x 10^6 /mL | FWB - Flow Wash Buffer (PBS + 2% FCS - Fetal Calf Serum) | 20 ul per well | Yes | 4.7 U/mL | No dilutions, only frozen and centrifuged at 10000 x g prior to use | Human Serum, Sigma-Aldrich, Product Number # H5667 | 2 | Own control: pooled positive PPS serum with confirmed reactivity in the CDC test. Selection criteria: Immunized patients in whom anti-HLA antibodies in class I and class II were detected. PRA = approx. 100%. | 5 positive control replicates in dilutions (neat, 1:2, 1:4, 1:8, 1:16) | 2 | 20 | Room temperature: approx. 23,5 C | Manual | FWB - Flow Wash Buffer (PBS + 2% FCS - Fetal Calf Serum) | 1 wash: 150 ul per well, 2-4 wash: 250 ul per well | 4 | Room temperature: approx. 23,5 C | Jackson ImmunoResearch, Code Number # 109-096-098 | Goat | FITC |  | 1:200 | PBS - Phosphate Buffered Saline | IgG-FITC in antibody cocktail = 0,175 uL per well | 5 | Room temperature: approx. 23,5 C | BD Biosciences, Catalog # 345766 | Yes |  | Manual | FWB - Flow Wash Buffer (PBS + 2% FCS - Fetal Calf Serum) | 1.4 | 2 | 23.5 | Yes | Other | PerCP | Yes | Yes | CD3 in antibody cocktail: 1,4 uL per well (with PBS, IgG-FITC, CD19) | 1 in 25 | 150 | 5 | 23.5 | BD Biosciences, Catalog #345789 | Yes |  | Manual | FWB - Flow Wash Buffer (PBS + 2% FCS - Fetal Calf Serum) | 0.7 | 2 | 23 | Phycoerythrin |  | Yes | No | Yes | CD19 in antibody cocktail: 0,7 uL per well (with PBS, IgG-FITC, CD3) | 1 in 50 | 150 | 5 | Room temperature: approx. 23,5 C | Manual | FWB - Flow Wash Buffer (PBS + 2% FCS - Fetal Calf Serum) | 1 wash: 150 ul per well, 2 wash: 250 ul per well | 2 | Room temperature: approx. 23,5 C | Other | FWB - Flow Wash Buffer (PBS + 2% FCS - Fetal Calf Serum) | 150 | Beckman Coulter DxFLEX | No | Yes | Alignment Check, Laser Check, Fluorescence Check | DxFLEX Daily QC Fluorospheres REF # C39283 | 30000 | 30000 | approx. 25000 | approx. 27000 | approx. 400 | approx. 1400 | Median |  | Log |  | Peak shift |  | More than negative control (MCF) + discriminator value | More than negative control (MCF) + discriminator value | T-Cells: 71 channel, B-Cells: 117 channel | T-Cells: 71 channel, B-Cells: 117 channel
Lab-1117 | PBMC isolation | 50µL | 250 000 WBC/tube | PBS | 50µL | Yes | 5mg/mL | undiluted + 1:3 (diluted with negative control) | pool of 10 different HLA antibody negative donors | 3 | HLA class 1 + HLA class 2 (invivogen) | 1 | 3 of undiluted sample and 3 of 1:3 diluted sample | 30 | cold on crashed ice | Manual | Cellwash (by Becton Dickinson) with 0,5% BSA (2-8°C) | 2mL | 2 | 6 | Dianova | Goat | Other | Alexa Fluor 488 | 1:1071 | Wash medium | 50 (antibody cocktail) | 15 | cold on crashed ice | Becton Dickinson | Yes |  | Manual | Cellwash (by Becton Dickinson) with 0,5% BSA (2-8°C) | 5 | 2 | 2 | Yes | Phycoerythrin |  | Yes | No |  |  | 2 | 15 | 6 | Becton Dickinson | Yes |  | Manual | Cellwash (by Becton Dickinson) with 0,5% BSA (2-8°C) | 2 | 2 | 6 | Other | APC | Yes | No | No |  |  | 2 | 15 | 2 | Manual | Cellwash (by Becton Dickinson) with 0,5% BSA (2-8°C) | 2 | 2 | 6 | Other | Wash medium | 100 | FACS Canto II - Becton Dickinson | No | No |  | CST Beads and 7-Color Setup Beads | [5 empty] | 2500 | Median |  | Log |  | Peak shift, Other | Ratio (mean sample / mean negative control) | >1,3 | >1,3 | same as for controls | Ratio (mean sample / mean negative control)
Lab-1159 | Received material was layered on Lymphodex and centrifuged with no brake. Collected mononuclear cell layer was washed 1xNaCl and 2 x PBS. Add pronase (incubation 13 min, 37 ℃), add DNAse (incubation 2 min, 23℃). Cell suspension in PBS + 2% FBS | 15 | 10*10^6 | PBS + 2%FBS | 30 | Yes | 500 ul pronase 4,7 U/ml for 10*10^6 cells | No | homemade (pooled anti-HLA negative sera from recipients) | 2 | homemade (pooled sera from hiperimmunized recipients) | 1 + dilutions (1:2, 1:4, 1:8, 1:16) | 2 | 20 | 23 | Manual | PBS + 2% FBS | 0,2 | 4 | 23 | Bio-rad | Goat | FITC |  | 200 | PBS | 0,25 | 5 | 23 | Becton Dickinson | Yes |  |  |  | 2 |  | 23 | Yes |  | PerCP | Yes | Yes | PBS | 2 in 50 |  | 5 |  | Becton Dickinson | Yes |  |  |  | 1 |  |  | Phycoerythrin |  | Yes | No | Yes | PBS | 1 in 50 |  | 5 | 23 | Manual | PBS + 2% FBS | 0,15 | 1 | 23 | Sheath fluid |  | 150 | BD FACS Lyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | CS&T Beads / FC Beads |  |  |  |  | 1000 |  | Median |  | Other | BiExpotential | Other | Median Fluorescence Intensity (MFI) Ratio | MFI PC / MFI NC >=2 | MFI PC / MFI NC >=2 | MFI sample / MFI NC >= 1,5 | none
Lab-1166 | Ficoll blood separation | Depends on cell concentration | 200000/tube | Phosphate Buffered Saline(PBS) | 100 (μl) | No |  | No | commercial control and also serum of volunteers male AB-blood group | 4 different negative controls | Pool of hypersensitive patients | 3 different dilutions | 3 triplicates tubes | 30' | Room Temperature(RT) | Manual | PBS | 2(ml) | 2 washes | RT | Jackson ImmunoResearch | Goat | Fab2 fragment |  | Depends on titration | PBS | 10(μl) | 30' | RT | CD3 (Beckman Coulter) | Yes |  | Manual | PBS | 10(μl) | 2 washes | RT | Yes | Phycoerythrin |  | Yes | No |  |  | 2(ml) | 30' | RT | CD19 (Beckman Coulter) | Yes |  | Manual | PBS | 10(μl) | 2 washes | RT | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  | 2 (ml) | 30' | RT | Manual | PBS | 2(ml) | 2 washes | RT | Phosphate Buffered Saline |  | 500(μl) | CytExpert for DxFlex(Beckman Coulter) | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | DxFLEX Daily Fluorospheres |  |  | 2500 | 5000 | 200 | 1000 | Median |  | Log |  | Median difference |  | Ratio>1,351 | Ratio>1,866
Lab-1179 | MACS prep | 25 | +- 2000 lymphocytes/μL | Terasaki park medium | 25 | No |  | No | In house: Serum from recalcified male AB plasma | 2 | In house: 3 sera with Bw4 antibodies + 2 sera with Bw6 antibodies | 2 | 2 | 20 | RT | Manual | 0.1% MACS BSA stock solution | 2.3 | 2 | RT | Abcam | Goat | FITC |  | 1/50 | PBS | 10 µL | 10 | RT | Analis (Beckman Coulter) | Yes | No |  |  | 10 |  | RT | Yes | Other | Allophycocyanine | Yes | No |  |  |  | 10 |  | Analis (Beckman Coulter) | Yes | No |  |  | 20 |  |  | Phycoerythrin |  | Yes | No | No |  |  |  | 10 | RT | Manual | 0.1% MACS BSA stock solution | 2.3 | 2 | RT | Other | 0.1% MACS BSA stock solution (wash buffer) | 300 | Navios 6/2 | No | Yes | Alignment Check, Laser Check | Flow Set beads, Flow check beads (Navios) |  |  | 2000 |  | 500 |  | Median |  | Linear |  | Median difference |  | Average median neg control + 2SD | Average median neg control + 2SD | Average median neg control + 2SD
Lab-1182 | [5 empty] | No | [24 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No | [18 empty] | No methods provided. Not tested.
Lab-1183 | negative sellection | 15 | 150000 | DPBS | 30 | Yes |  | no | inhouse | 3 | Inhouse | 2 | 2 | 20 | 22 | Manual | DPBS+BFS | 0.2 | 4 | 22 |  | Goat | [8 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [5 empty] | Manual | DPBS+BFS | 0.2 | 2 | 22 | Phosphate Buffered Saline |  |  |  | Yes | Yes | [6 empty] | 500 |  | Other
Lab-1197 | Ficoll | 50 | 0,2x10^6 for 50uL | RPMI | 50 | No |  | NO | NIBSC | 2 | NIBSC | 2 | 2 | 30 | 22 | Manual | PBS BSA 1% | 0,2 | 4 | 22 | Jackson immunoresearch | Goat | FITC |  | 1/50 | PBS BSA 1% | 20 | 10 | 22 | Beckman Coulter | Yes |  |  |  | 4 |  | 22 | Yes | Phycoerythrin |  | Yes | No |  |  |  | 10 |  | Beckman Coulter | Yes |  |  |  | 2 |  |  | Phycoerythrin-cyanine 5 |  | Yes | No | Yes |  |  |  | 10 | 22 | Manual | PBS BSA 0,5% | 0,6 | 3 | 22 | Phosphate Buffered Saline |  | 150 | FACSFLOW | Yes | Yes | Laser Check |  | 20000 | [5 empty] | Mean |  | Log |  | Other | Ratio | ratio > 2,3 | ratio > 1,7
Lab-1248 | IMMUNOMAGNETIC CELL SEPARATION | 40 | 3.0 X 10^6 CELL/ ML | DPBS + 2% FBS | 50 | No |  | No | 1 POOLED MALE AB DONOR AND LS-NC | 2 | POOL OF HIGH PRA PATIENTS'S SERA | 1 | 2 | 20 | 24 | Manual | DPBS + 2% FBS & DPBS | 4 | 2 | 24 | JACKSON IMMUNORESEARH LABORATORIES | Goat | FITC |  | NIL | ddH2O | 8 | 15 | 0 | CLONE SK7, BD | No | Yes | Manual | DPBS WITH 2% FBS AND DPBS | 10 | 1 | 0 | No | Phycoerythrin |  | Yes | No |  |  | 4 | 15 | 24 | NIL | No | No |  | NIL | NIL | NIL | NIL |  |  | No | No | No | NIL | NIL | NIL | NIL | NIL | Manual | DPBS WITH 2% FBS AND DPBS | 4 | 1 | 24 | Other | STAIN BUFFER | 600 | FACSLYRIC | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | CS&T BEADS |  |  | 3000 | 10000 |  |  | Median |  | Log |  | Peak shift, Median difference, % shift |  | >4500 (ANTI IgG FITC) |  | <1086 (ANTI IgG FITC VALUE) | >257 (ANTI IgG FITC SHIFT) WITH RATIO >2.0 | Interpretation was based on combined assessment of shift, Anti-IgG FITC intensity and ratio values; shift alone was not used to define positivity.
Lab-1258 | [5 empty] | No | [24 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No
Lab-1277 | Magnetic Beads | 15 µL / well | 1.70.10E6 cell/ml | PBS | 30 | Yes | 4.7 kuntz Units/mL | no | Home made SAB pool | 2 | SFHI Single Ag class I positive control | 2 | 2 | 20 | 22°C | Manual | PBS 2% SVF | 0.2 | 4 | 22°C | Jackson | Goat | FITC |  | 2.5 µL / mL | PBS | 0.125 µL / well | 5 | 22°C | Becton Dickinson CD3 PerCP Cy5.5 | Yes |  | Manual | PBS 2% SVF | 2 µL / well | 2 | 22 | Yes | Phycoerythrin-cyanine 5 |  | Yes | Yes | PBS | 40 µL /mL | 0.2 | 5 | 22 | Becton Dickinson CD19 PE | Yes |  | Manual | PBS 2% SVF | 1 µL / well | 2 | 22 | Phycoerythrin |  | Yes | No | Yes | PBS | 20µL / mL | 0.2 | 5 | 22 | Manual | PBS 2%SVF | 0.2 | 2 | 22 | Other | PBS 2% SVF | 200 | FACs VIA BD | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | CST Beads |  | 20000 | 5000 |  | 200 |  | Median |  | Log |  | Other | Fluorescence Ratio between sample and negative control | 1.8 | 1.9 | ratio | ?
Lab-1283 | [5 empty] | No | [24 empty] | No | [6 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No