## 2B-3-4-2026 G3 Methods
Lab Code | Cell Preparation Method | Cell Volume/Tube | Cell Concentration | Cell Diluent | Serum Volume/Tube | Pronase Treatment | Pronase Treatment Details | Serum Dilutions | Negative Control | Number of Negative Replicates | Positive Control | Number of Positive Replicates | Number of Crossmatch Test Replicates | Incubation Time (Mins) | Sensitisation Temperature | First Wash Step | Wash Medium | First Wash Volume/Tube | Number of Washes | First Wash Temperature | Manufacturer | Anti-human IgG Reagent Raised In | Anti-human IgG Labelled With | Detection Stage More Details | Detection Stage Dilution | Detection Stage Diluted In | Detection Stage Volume / Tube | Detection Stage Incubation Time (mins) | Detection Stage Incubation Temperature | Anti-T Cell Reagent Manufacturer | Anti-Human IgG Reagent added with Anti-T cell | Is there a wash step before adding Reagent | Anti-T Cell Reagent Wash Step Details | Anti B-Cell Reagent Wash Step Details | Anti-T Cell Reagent Wash Medium | T-Cell Volume/Tube | Anti-T Cell Reagent Number of Washes | T-Cell Incubation Temp | Is Anti T cell reagent with Anti-B cell reagent | Anti-T Cell Reagent Labelled with | Anti-T Cell Reagent Other Details | Anti-T Cell Reagent CD3 | Anti-T Cell Reagent Diluted | Anti-T Cell Reagent Diluted In | Anti-T Cell Reagent Dilution | Anti-T Cell Reagent Volume/Tube | T-Cell Incubation Time (mins) | Anti-T Cell Reagent Temperature | Anti B-Cell Reagent Manufacturer | Is Anti B Cell Reagent added with Anti human IgG | Is there a wash step before adding Anti B Cell | Anti B-Cell Reagent Wash Step Details | Anti B-Cell Reagent Wash Medium | B-Cell Volume/Tube | Anti B-Cell Reagent Number of Washes | Anti B-Cell Reagent Temperature | Anti B-Cell Reagent Labelled With | Anti B-Cell Reagent Other Details | Anti B-Cell Reagent CD19 | Anti B-Cell Reagent CD20 | Anti B-Cell Reagent Diluted | Anti B-Cell Reagent Diluted In | Anti B-Cell Reagent Dilution | Anti B-Cell Reagent Volume/Tube | B-Cell Incubation Time (mins) | B-Cell Incubation Temp | Final Wash Step Details | Final Wash Medium | Final Wash Volume/Tube (µl) | Final Wash Number of Washes | Final Wash Temperature (C) | Final Wash Resuspension of Cells | Final Wash Other Details | Final Wash Volume/Tube (ml) | Data Acquisition Make of Flow Cytometer | Data Acquisition Live Gated | Data Acquisition Calibration Beads Used | Data Acquisition Used For | Data Acquisition Type of Beads | Total Events Minimum | Total Events Maximum | T-Cells Minimum | T-Cells Maximum | B-Cells Minimum | B-Cells Maximum | Data Acquisition Parameter Assess Peak Performance | Data Acquisition Peak Performance Please Specify | Data Acquisition Scale Used | Data Acquisition Scale Used Please Specify | Data Analysis Parameter used to Assess Difference | Data Analysis Other Details | Data Analysis T Cell Miniumum Value for PCS | Data Analysis B Cell Minimum Value for PCS | Data Analysis Positive From Negative | Data Analysis Distinguish Binding of Human IgG | Comments
Lab-1038 |  | 20 | 300000 | PBS | 20 | No |  | No | Serum from AB-, non-transfused males | 2 | commercial (anticlass I and class II, One Lambda) | 2 | 2 | [8 empty] | Goat | FITC | [7 empty] | No | [7 empty] | Yes | Phycoerythrin |  | Yes | No | [6 empty] | No | [6 empty] | Phycoerythrin-cyanine 5 |  | Yes | No | No | [13 empty] | FACS Lyric (BD Biosciences) | No | No |  |  |  |  | 2000 |  | 2000 |  | Median |  | Log |  |  | Ratio Median (sample)/Median (negative control)
Lab-1041 | isolation of PBMC from whole blood | 200 | minimum 1x106 | PBS | 10 | Yes | 1 mg¦ml | none | negative AB-sera local | 2 | pooled local sera | 2 | 2 | 15 | room temperature | Manual | PBS for T-cells, PBS with 1% fetal calf sera for B-cells | 0.7 | 4x B- and T-cells | room temperature | T-cells BD Biosciences | Goat | Fab2 fragment |  | 1:100 | PBS | 200 | 15 | room temperature | BD Biosciences | No | Yes | Manual | Manual | PBS | 200 | 2 | room temperature | No | Phycoerythrin |  | Yes | Yes | PBS | 1 in 20 | 0.7 | 15 | room temperature | DAKO | Yes | No | Manual | PBS 1% fetal calf sera | 0.7 | 2 | room temperature | Phycoerythrin |  | Yes | No | Yes | PBS | 1 in 20 | 0,7 | 15 | room temperature | Manual | PBS | 400 | 1 | room temperature | Phosphate Buffered Saline |  | 0.7 | BD Biosciences FACSLyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check | C,S and T-Beads |  |  | 1000 | 5000 | 20 | 1000 | Median |  | Linear |  | Median difference |  | minimum 150 MFI between negative and positve control | minimum 100 MFI between negative and positve control | median negative control + 60 MFI for T cells, median negative control + 18 MFI for B cells | median negative control + 60 MFI for T cells, median negative control + 18 MFI for B cells
Lab-1059 | isolation of lymphocytes by gradient centrifugation | 50 |  |  | 50 | No |  | no | NIBSC | 2 | Local positive control | 2 | 2 | 30 | 22 | Manual | Cell wash+FCS | 1 | 3 | 22 |  | Goat | Fab2 fragment |  | no |  | 1.0 | 30 | 22 | BD | Yes |  | Manual | Manual | Cell wash with FCS | 50 | 3 | 22 | Yes | Phycoerythrin-cyanine 5 |  | Yes | No |  |  | 1.0 | 30 | 22 | BD | Yes |  | Manual | Cell wash with FCS | 50 | 3 | 22 | Phycoerythrin |  | Yes | No | No |  |  | 1.0 | 30 | 22 | Manual | cell wash with FCS | 350 | 2 | 22 | Fixative |  | 1.0 | Facs Canto | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification |  | 5000 | 50000 | 3 000 | 50 000 | 500 | 1000 | Other | Geo mean | Linear |  | Other | Geo mean linear value | 3 SD | 3SD | 3SD | 3SD
Lab-1068 | Immunomagnetic negative selection | 30 | 8000 cells/microL | PBS + 2 % fetal calf serum + EDTA 1mM | 25 | No |  | No | NIBSC | 2 | NIBSC + in-house material | 1 for each preparation (NIBSC and in-house) | 2 | 30 | 4 | Manual | PBS + 2 % fetal calf serum + EDTA 1mM | 0,4 | 3 | 4 | Jackson ImmunoResearch Laboratories, Inc. 109-096-097 | Goat | FITC |  | Reconstituted with 750 microL dH2O | PBS + 2 % fetal calf serum + EDTA 1mM | 20 | 30 | 4 | BD Biosciences™ 345766 | Yes |  | Manual | Manual | PBS + 2 % fetal calf serum + EDTA 1mM | 20 | 2 | 4 | Yes | Other | PerCP | Yes | Yes | PBS + 2 % fetal calf serum + EDTA 1mM | 1 in 4 | 0,4 | 30 | 4 | BD Biosciences™ 345789 | Yes |  | Manual | PBS + 2 % fetal calf serum + EDTA 1mM | 20 | 2 | 4 | Phycoerythrin |  | Yes | No | Yes | PBS + 2 % fetal calf serum + EDTA 1mM | 1 in 4 | 0,4 | 30 | 4 | Manual | PBS + 2 % fetal calf serum + EDTA 1mM | 300 | Because we incubate anti-IgG, -CD3 and -CD19 simultaneously, the final wash is the same as previous wash steps indicate for each reagent | 4 | Other | PBS + 2 % fetal calf serum + EDTA 1mM | 0,4 | BD Lyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | BD CS&T Beads 656504 |  |  | 500 |  | 500 |  | Median |  | Log |  | % shift |  | 50 % increase in shift compared with negative control | 70 % increase in shift compared with negative control | 50 % shift for T-cells and 70 % shift for B-cells | 50 % shift for T-cells and 70 % shift for B-cells
Lab-1072 | Stemcell B Isolation Kit | 3 |  |  | 3 | No |  | No | [21 empty] | Stemcell | No | [7 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No
Lab-1074 | [5 empty] | No |  | yes 1/2 | local | 2 | local | 1 | 2 | 30 | 25 | Manual |  | 1 | 2 | 25 |  | Rabbit | FITC |  | 1/10 |  | 25 | 40 | 4 | Beckman coulter | Yes |  | Manual | Manual |  | 2,5 | 2 | 4 | Yes |  | PERCP | Yes | No |  |  | 1 | 40 | 25 | Beckman coulter | Yes |  | Manual |  | 2,5 | 2 | 25 | Phycoerythrin |  | Yes | No | No |  |  | 1 | 40 | 4 | Manual |  | 500 | 1 | 25 | Phosphate Buffered Saline |  | 1 | FACSCANTO II | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification |  |  |  | 10.000 | 100.000 | 1.000 | 2000 | Median |  | Log |  | Median difference |  | 53 | 167
Lab-1100 | flow citometry | 200 | 10.000.000 | PBS | 15 | No |  | no | AB serum | 2 | pool of hyperimmunized sera | 2 | 2 | 20 | 22 | Manual | PBS 2% FBS | 200 | 1 | 4 | Agylent | Rabbit | FITC |  | 1:10 | PBS | 40 | 10 | 22 | BD | Yes |  | Manual | Manual | PBS 2%FCS | 40 | 1 | 22 | Yes | Other | Per CP | Yes | Yes | PBS | 1:10 | 200 | 10 | 4 | BD | Yes |  | Manual | PB 2% FCS | 200 | 1 | 4 | Other | APC | Yes | No | Yes | PBS | 1:20 | 200 | 10 | 22 | Manual | PBS 2% FCS | 200 | 1 | 4 | Other | PBS 2% BSA | 200 | BD FACS Lyric | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | CST beads | 1000 | 3 minutes | 1000 | 3 minutes | 1000 | 3 minutes | Median |  | Log |  | Median difference
Lab-1115 | isolated lymphocyte samples distributed (no prepatation requirement) | 50 | 1 x 10^6 cells/100ul | Phosphate Buffered Saline (PBS) | 50 | No |  | No | NS Negative control (One Lambda, Inc) | 2 | inno-train | 2 | 2 | 30 | RT, 20-25 | Manual | Phosphate Buffered Saline (PBS) | 4 | 2 | 22 | Jackson ImmunoResearch Laboratories, Inc. | Goat | FITC |  | 1:50 | PBS | 50 | 30 | 20-25 | BECKMAN COULTER | Yes |  |  |  | PBS | 10 | 2 | 2-8 | Yes | Phycoerythrin |  | Yes | No |  |  | 4 | 20 | 22 | BECKMAN COULTER | Yes |  |  | PBS | 10 | 2 | 22 | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  | 4 | 20 | 2-8 | Manual | PBS | 500 | 1 | 22 | Phosphate Buffered Saline |  | 4 |  | Yes | No |  |  | 5000 | 20000 | 4500 | 12000 | 450 | 2000 | Median |  | Log |  | Median difference |  | 50 | 80 | 1.5
Lab-1127 | [5 empty] | No | [24 empty] | No | [7 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No
Lab-1131 | [5 empty] | No | [24 empty] | No | [7 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [5 empty] | Manual | [8 empty] | No | No | [12 empty] | Other | ratio
Lab-1135 | Cell Washing+Cell Freezing on 23.07.2026, Cell Thawing for FACS Assay on 28.07.2026 | 25 | 300.000 per well | RBMI-1640 | 25 | Yes | 2mg/ml | our Positive Control: 1:50 | in-house | 2 | in-house | 2 | 2 | 30 | 4°C | Manual | FACS Buffer (500ml DPBS + 2,5g BSA + 1ml Natriumazid (NaN3) | 0,125 | 2 | 4°C | Jackson ImmunoResearch | Goat | FITC |  | 1:100 | FACS Buffer (500ml DPBS + 2,5g BSA + 1ml Natriumazid (NaN3) | 50 | 30 | 4 | BioLegend | Yes |  | Manual | Manual | FACS Buffer (500ml DPBS + 2,5g BSA + 1ml Natriumazid (NaN3) | 5 | 2 | 4 | Yes | Other | BV605 | Yes | No |  |  | 0,125 | 30 | 4 | BioLegend | Yes |  | Manual | FACS Buffer (500ml DPBS + 2,5g BSA + 1ml Natriumazid (NaN3) | 1 | 2 | 4 | Other | APC | Yes | No | No |  |  | 0,125 | 30 | 4 | Manual | FACS Buffer (500ml DPBS + 2,5g BSA + 1ml Natriumazid (NaN3) | 300 | 2 | 4 | Other | FACS Buffer (500ml DPBS + 2,5g BSA + 1ml Natriumazid (NaN3) | 0,125 | BD LRS Fortessa | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | BD CS&T Beads | 15000 |  | 10000 |  | 5000 |  | Other | GeoMean | Linear |  | Mean difference |  | 1,5 times higher than negative control sample | 1,7 times higher than negative control sample | -- | --
Lab-1152 |  | 25 | 2*10^7 cells/ml | PBS | 30 | Yes | 0,5 mg/ml | NO | Commercial Human AB Serum | 1 | Patients | 2 | 1 | 30 | 4 | Automated cell washer | PBS | 0.8 | 2 | RT | Invitrogen | Goat | FITC |  | 1:80 | PBS/FBS | 80 | 30 | 4 | Beckman Coulter | No | No | Automated cell washer | Automated cell washer | PBS | 5 | 2 | 4 | Yes | Phycoerythrin |  | Yes | No |  |  | 0.8 | 20 | RT | Beckman Coulter | No | No | Automated cell washer | PBS | 5 | 2 | RT | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  | 0.8 | 20 | 4 | Automated cell washer | PBS | 350 | 2 | RT | Phosphate Buffered Saline |  | 0.8 | Beckman Coulter | Yes | Yes | Alignment Check | Polystyrene | 10000 | 25000 | 10000 | 20000 | 1000 | 1500 | Median |  | Log |  | Other | Median Ratio | 1.3 ratio | 1.3 ratio | 1.2
Lab-1154 | cell lysis | 100 |  |  | 50 | No |  |  | own | 2 | own |  | 2 | 30 | RT | Manual | Washing buffer | 2 | 3 | RT | Jackson Immuno research | Goat | FITC |  | 1/80 | wash buffer | 25 | 20 | RT | BD | Yes |  |  |  | wash buffer | 25 | 3 | on ice | No | Phycoerythrin |  | Yes | Yes | washbuffer | 1/2 | 2 | 20 | RT | BD | Yes |  |  | wash buffer | 25 | 3 | RT | Phycoerythrin |  | Yes | No | Yes | washbuffer | 1/2 | 2 | 20 | on ice | Manual | washbuffer | 300 | 3 | RT | Fixative |  | 2 | Beckman Coulter DxFlex | No | No |  |  |  |  | 1000 | 2500 | 1000 | 2500 | Median |  | Linear, Log |  | Other | Ratio of Median fluorescence of sample vs negative control. | N.A. | N.A. | Ratio 1.5 for T cells and 2 for B-cells. | Ratio 1.5 for T cells and 2 for B-cells.
Lab-1169 | Not applicable | 50 |  |  | 50 | No |  | No | Inno-train Diagnostik GmbH | One | Inno-train Diagnostik GmbH | One | One | 30 min | RT | Manual | PBS | 2 | 2 | RT | Jackson ImmunoResearch INC | Goat | FITC |  | 1 in 50 | PBS | 50 | 20 min | RT | Beckton Dickinson | Yes |  |  |  |  | 10 |  | RT | No | Phycoerythrin |  | Yes | No |  |  |  | 20 min | RT | Immunotech Beckman Coulter | Yes |  |  |  | 10 |  |  | Phycoerythrin |  | Yes | No | No |  |  |  | 20 min | RT | Manual | PBS | 500 | One | RT | Phosphate Buffered Saline |  | 2 | Navios EX Beckman Coulter | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check | Flow Check Pro Fluorospheres by Beckman Coulter |  |  | 5000 |  | 500 |  | Median |  | Log |  | Median difference |  |  |  | ratio med MFI TEST/MFI neg > mean +2SD
Lab-1184 | To | 15 | 10 | DPBS | 30 | No |  |  | Human AB serum | 3 | Pool Hiper sensititized | 2 | 2 | 20 | 22 | Manual | DPBS | 150 | 3 | 22 |  | Goat | FITC |  | 1/20 | DPBS | 20 | 15 | 22 | Coulter | No |  |  | Manual |  |  |  |  | No |  |  | No | No | [5 empty] | BD | No | No | Manual | DPBS | 4 | 3 | 22 | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  | 4 | 15 | 22 | Manual | DPBS | 150 | 3 | 22 | Phosphate Buffered Saline |  | 150 | FACS Lyric | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | CST | 30000 | 50000 | 5000 | 30000 | 500 | 1000 | Median |  |  |  | Peak shift, Median difference |  | 260 | 1190 | 1.2 | 1.5
Lab-1207 | Washed in PBS, treated with DNAse and Pronase | 100 | 1,000,000 per mL | PBS with 0.1% Bovine Serum Albumin | 100 | Yes | 32 U/mL | No | In-house serum pool from male AB blood donors tested negative for anti-HLA antibodies using Luminex testing | 1 | In-house serum pool from immunized patients with known anti-HLA antibodies | 1 | 1 | 30 | 21 | Manual | PBS with 0.1% Bovine Serum Albumin | 2 | 3 | 21 | DAKO/Agilent | Rabbit | FITC |  | 1:100 | PBS with 0.1% Bovine Serum Albumin | 100 | 20 | 21 | Waters Biosciences | No | Yes | Manual | Manual | PBS with 0.1% Bovine Serum Albumin | 2 | 1 | 21 | Yes | Phycoerythrin |  | Yes | No |  |  | 2 | 20 | 21 | BioLegend | No | Yes | Manual | PBS with 0.1% Bovine Serum Albumin | 2 | 1 | 21 | Other | BrilliantViolet785 | No | Yes | No |  |  | 2 | 20 | 21 | Manual | PBS with 0.1% Bovine Serum Albumin | 300 | 1 | 21 | Other | PBS with 0.1% Bovine Serum Albumin | 2 | Agilent/ACEA NovoCyte 3000 | No | Yes | Alignment Check, Laser Check, Fluorescence Check, Fluorescence Quantification | 6-peak Rainbow Beads (Waters biosciences), NovoCyte QC Calibration Beads (Agilent), Compensation Beads (Waters biosciences) |  |  | 400 |  | 400 |  | Median |  | Log |  | Other | MFI-units are converted to MESF-units based on Rainbow calibration beads. Readout is Log(MESF sample / MESF negative) | 0.30 | 0.30 | 0.25 | 0.30 | Values below 0.25 are reported as Negative. Values between 0.25 and 0.30 are reported as Equivocal. Values above 0.30 are reported as Positive
Lab-1208 | Washed in PBS/0.1% BSA twice before used | 10 | N/A | PBS/0.1% BSA | 30 | No |  | No - serum used as supplied | Heat-inactivated pooled serum from male AB donors (Sigma Aldrich) | 2 | Serum mix from five highly immunized patients | 2 | 2 | 15 | Room temperature | Manual | PBS/0.1% BSA | 0.25 | 3 | Room temperature | Jackson Immuno Research | Goat | FITC |  | 1:200 | PBS/0.1% BSA | N/A | 10 | Room temperature | BD Biosciences | Yes |  | Manual | Manual | PBS/0.1% BSA | 10 | 1 | Room temperature | Yes | Other | APC | Yes | Yes | PBS/0.1% BSA | 1 in 10 | 0.15 | 10 | Room temperature | BD Biosciences | Yes |  | Manual | PBS/0.1% BSA | 10 | 1 | Room temperature | Phycoerythrin |  | Yes | No | Yes | PBS/0.1% BSA | 1 in 10 | 0.15 | 10 | Room temperature | Manual | PBS/0.1% BSA | 250 | 1 | Room temperature | Phosphate Buffered Saline |  | 0.2 | BC Navios EX | Yes | Yes |  | Flow Check and Flow Set | [6 empty] | Median |  | Log |  | Median difference |  | >10x negative control median fluorescence intensity | >10x negative control median fluorescence intensity | >1.5x negative control median fluorescence intensity | >1.5x negative control median fluorescence intensity
Lab-1235 | fresh drawn heparin blood, dilute 1:1 in PBS | 25 | 20x10E6/ml | PBS/BSA | 25 | No | NA | no | commercial | 3 | commercial | 1 | 2 | 30 | RT | Manual | PBS | 2ml | 3 | RT | Dako | Goat | FITC |  | 1:100 | PBS/BSA | 50 | 30 | RT | Beckman Coulter | No | [7 empty] | Yes | Other | pacific blue | Yes | No | [6 empty] | Yes | [8 empty] | Yes | No | No | [5 empty] | Manual | PBS |  | 3 | RT | Phosphate Buffered Saline |  | 2ml | DxFlex Beckman Coulter | Yes | No | [8 empty] | Mean |  | Log |  | Peak shift |  | ratio 1.39 | ratio 1.95 | ratio pos to neg | gating on naive (sIgG neg) B-cells
Lab-1256 | Lymphocyte Separation Medium™ (density gradient medium) | 25 | 350000 cells | PBS | 50 | Yes | 0.5 mg/ml | no | Patients Pool | 2 | Patients Pool | 2 | 0 | 20 | 20-25 | Automated cell washer | PBS | 0.8 | 2 | 20-25 | Jackson ImmunoResearch | Goat | FITC |  | 1/180 | PBS | 80 | 10 | 20-25 | Beckman Coulter | Yes |  | Automated cell washer | Automated cell washer | PBS | 10 | 2 | 20-25 | Yes | Phycoerythrin |  | Yes | No |  |  | 0.8 | 10 | 20-25 | Beckman Coulter | Yes |  | Automated cell washer | PBS | 10 | 2 | 20-25 | Phycoerythrin-cyanine 5 |  | Yes | No | No |  |  | 0.8 | 10 | 20-25 | Automated cell washer | PBS 1X | 350 | 2 | 20-25 | Phosphate Buffered Saline |  | 0.8 | Beckman Coulter | Yes | Yes | Alignment Check, Laser Check, Fluorescence Check | Flow Check |  |  |  |  | 1000 |  | Median |  | Log |  | Other | Ratio | Ratio > 1.8 | Ratio >2.0 | Ratio | Ratio
Lab-1305 | [5 empty] | No |  |  | AB sera | 2 | pool sera | 1 | 1 | [8 empty] | Goat | FITC | [7 empty] | No | [7 empty] | No |  |  | No | No | [6 empty] | No | [8 empty] | No | No | No | [14 empty] | No | No | Fluorescence Quantification | [7 empty] | Mean, Median