## 2B-3-2026 T cells (G3)
Result summary | Serum 1 | Serum 2 | Serum 3 | Serum 4
Total distributed | 20 | 20 | 20 | 20
Total submitted | 20 | 20 | 20 | 20
Not Tested | 2 | 3 | 2 | 2
Positive | 8 | 13 | 0 | 10
Negative | 10 | 4 | 18 | 8
Equivocal | 0 | 0 | 0 | 0
% Positive | 44.44% | 76.47% | 0% | 55.56%
% Negative | 55.56% | 23.53% | 100% | 44.44%
% Equivocal | 0% | 0% | 0% | 0%
Consensus | No consensus | Positive | Negative | No consensus
Unacceptable | N/A | Negative, Equivocal | Positive, Equivocal | N/A
Number acceptable | N/A | 13 | 18 | N/A
Number unacceptable | N/A | 4 | 0 | N/A
Assessed | No | Yes | Yes | No
Laboratory Code | T-cells Serum 1 | T-cells Serum 2 | T-cells Serum 3 | T-cells Serum 4 | T-cells Serum 1 Cytometer Reading | T-cells Serum 2 Cytometer Reading | T-cells Serum 3 Cytometer Reading | T-cells Serum 4 Cytometer Reading | Units for cytometer readings | Cell Viability | Comments | T-cells Local Negative Control | T-cells Local Positive Control | T-cells Positive Result Value | Date received | Date Tested | Submitted | Distribution Year | Reason for not testing | Confirm Reason | Tested | Correct | Assessed | Tcell
Lab-1038 | Positive | Positive | Negative | Positive | 154 | 297 | 68 | 198 | fluorescence medians | 90 |  | 84 | 4261 | Ratio sample/negative control > 1.5 | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1041 | Not Tested | Not Tested | Not Tested | Not Tested | [6 empty] | No results, due the late delivery. |  |  |  | 2026-07-23 |  | Yes | 2026 | Sample delivery delay | No results, due late delivery. | 0 | 0 | 0 | Yes
Lab-1059 | Positive | Positive | Negative | Positive | 19.0 | 26.0 | 10.0 | 23.0 | Geo mean linear value | 85 |  | 10.0 | 93.0 | 3SD | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1068 | Negative | Positive | Negative | Positive | 1,39 | 2,15 | 0,74 | 1,67 | % shift |  |  | 1,0 | 1,84 | 50% increase in shift compared to negative control | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1072 | Not Tested | Not Tested | Not Tested | Not Tested | [5 empty] | 95 |  |  |  |  | 2026-07-23 | 2026-07-24 | Yes | 2026 | Equivocal | equivoical | 0 | 0 | 0 | Yes
Lab-1074 | Positive | Not Tested | Negative | Positive | 278 |  | 145 | 267 | median log channel | 85 |  | 122 | 7130 | 53 median log channel sihft | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 3 | 1 | 1 | Yes
Lab-1100 | Negative | Negative | Negative | Negative | 187 | 195 | 165 | 192 | median log channel | 99 |  | 180 | 4472 | NEG + 2SD (207) | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 1 | 2 | Yes
Lab-1115 | Negative | Negative | Negative | Negative | [5 empty] | 89 |  |  |  |  | 2026-07-23 | 2026-07-23 | Yes | 2026 | Cell viability too low | cell viability too low | 4 | 1 | 2 | Yes
Lab-1127 | Negative | Positive | Negative | Negative | 1.3 | 1.78 | 0.85 | 1.28 | MFI | 97.7 |  | - | 5.5 | >1.75 | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1131 | Positive | Positive | Negative | Positive | 387,2 | 660,3 | 147,0 | 409,7 | Median Log Channel | 89 |  | 200,4 | 18663,0 | 2SD | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1135 | Negative | Positive | Negative | Negative | 432 | 471 | 173 | 413 |  | 95 |  | 296 | 9531 | 1,5 x higher than negative control | 2026-07-23 | 2026-07-28 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1152 | Negative | Negative | Negative | Negative | 0.665 | 0.715 | 0.531 | 0.715 |  |  |  | 0.536 | 2.930 | 3SD | 2026-07-27 | 2026-07-27 | Yes | 2026 |  |  | 4 | 1 | 2 | Yes
Lab-1154 | Positive | Positive | Negative | Positive | 12622 | 20611 | 6280 | 20328 |  |  |  | 6603 | 23627 | A ratio of 1,5 as compared to the average for the negative control replicates | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1169 | Negative | Negative | Negative | Negative | 0.464 | 0.499 | 0.342 | 0.494 | median log fluorescence intensity | 90 |  | 0.384 | 14.9 | 2SD | 2026-07-23 | 2026-07-27 | Yes | 2026 |  |  | 4 | 1 | 2 | Yes
Lab-1184 | Positive | Positive | Negative | Positive | 367 | 462 | 232 | 919 | median channel | 95 |  | 263 | 1400 | >1.2 (263) | 2026-07-23 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1207 | Positive | Positive | Negative | Positive | 0.324 | 0.550 | -0.127 | 0.409 |  | 0 | Viability not assessed | 0 | 1.228 | 0.30 | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1208 | Negative | Positive | Negative | Negative | 5.7 | 9.7 | 3.6 | 6.4 | Median fluorescence intensity | 97.6 |  | 4.5 | 119 | >1.5X negative control median flourescence intensity | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1235 | Negative | Positive | Negative | Negative | ratio 1.09 | ratio 1.46 | ratio 0.98 | ratio 1.15 | MFI | 98 |  | MFI 8609 | MFI 75629 | ratio >1.39 | 2026-07-21 | 2026-07-22 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1256 | Negative | Positive | Negative | Positive | 0.531 | 0.734 | 0.273 | 0.715 | Median Log Channel | 90 |  | 0.360 | 6.41 | Ratio >1.8 | 2026-07-24 | 2026-07-27 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1305 | Positive | Positive | Negative | Positive | [10 empty] | 2026-07-23 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes

## 2B-4-2026 T cells (G3)
Result summary | Serum 1 | Serum 2 | Serum 3 | Serum 4
Total distributed | 20 | 20 | 20 | 20
Total submitted | 20 | 20 | 20 | 20
Not Tested | 3 | 1 | 1 | 2
Positive | 9 | 0 | 0 | 7
Negative | 8 | 19 | 19 | 11
Equivocal | 0 | 0 | 0 | 0
% Positive | 52.94% | 0% | 0% | 38.89%
% Negative | 47.06% | 100% | 100% | 61.11%
% Equivocal | 0% | 0% | 0% | 0%
Consensus | No consensus | Negative | Negative | No consensus
Unacceptable | N/A | Positive, Equivocal | Positive, Equivocal | N/A
Number acceptable | N/A | 19 | 19 | N/A
Number unacceptable | N/A | 0 | 0 | N/A
Assessed | No | Yes | Yes | No
Laboratory Code | T-cells Serum 1 | T-cells Serum 2 | T-cells Serum 3 | T-cells Serum 4 | T-cells Serum 1 Cytometer Reading | T-cells Serum 2 Cytometer Reading | T-cells Serum 3 Cytometer Reading | T-cells Serum 4 Cytometer Reading | Units for cytometer readings | Cell Viability | Comments | T-cells Local Negative Control | T-cells Local Positive Control | T-cells Positive Result Value | Date received | Date Tested | Submitted | Distribution Year | Reason for not testing | Confirm Reason | Tested | Correct | Assessed | Tcell
Lab-1038 | Positive | Negative | Negative | Positive | 293 | 97 | 88 | 213 | fluorescence medians | 90 |  | 119 | 2053 | Ratio sample/negative control > 1.5 | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1041 | Not Tested | Not Tested | Not Tested | Not Tested | [6 empty] | No results, due the late delivery. |  |  |  | 2026-07-23 |  | Yes | 2026 | Sample delivery delay | No results, due the late delivery. | 0 | 0 | 0 | Yes
Lab-1059 | Positive | Negative | Negative | Positive | 21.0 | 10.0 | 11.0 | 17.0 | Geo mean linear value | 80 |  | 10.0 | 137.0 | 3SD | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1068 | Positive | Negative | Negative | Positive | 2,39 | 1,14 | 0,88 | 1,79 | % shift |  |  | 1,0 | 3,61 | 50% increase in shift compared to negative control | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1072 | Negative | Negative | Negative | Negative | 1383.00 | 1388.35 | 1202.65 | 1294.25 | MFI | 94 |  | 1326.30 | 5682.60 | 500 log channel shift | 2026-07-23 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1074 | Not Tested | Negative | Negative | Negative |  | 132 | 118 | 197 | median log channel | 85 |  | 152 | 2110 |  | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 3 | 2 | 2 | Yes
Lab-1100 | Negative | Negative | Negative | Negative | 207 | 182 | 180 | 199 | median log channel | 98 |  | 192 | 3383 | NEG + 2SD (219) | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1115 | Negative | Negative | Negative | Negative | [5 empty] | 92 |  |  |  |  | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1127 | Negative | Negative | Negative | Negative | 1.15 | 0.89 | 0.80 | 1.11 | MFI | 97.6 |  | - | 1.80 | >1.75 | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1131 | Positive | Negative | Negative | Positive | 327,2 | 166,1 | 149,1 | 282,2 | Median Log Channel | 87 |  | 160,8 | 38688,5 | 2SD | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1135 | Positive | Negative | Negative | Negative | 381 | 151 | 135 | 208 |  | 95 |  | 230 | 26624 | 1,5 x higher than negative control | 2026-07-23 | 2026-07-28 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1152 | Negative | Negative | Negative | Negative | 0.671 | 0.613 | 0.597 | 0.636 |  |  |  | 0.613 | 0.789 | 3SD | 2026-07-27 | 2026-07-27 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1154 | Positive | Negative | Negative | Positive | 11099 | 5330 | 5110 | 7960 |  |  |  | 5337 | 7840 | A 1,5 ratio between the serum and the negative control replicates | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1169 | Negative | Negative | Negative | Negative | 0.398 | 0.271 | 0.283 | 0.327 | median log fluorescence intensity | 90 |  | 0.288 | 9.44 | 2SD | 2026-07-23 | 2026-07-27 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1184 | Positive | Negative | Negative | Positive | 403 | 237 | 300 | 648 | median channel | 95 |  | 310 | 986 | >1.2 (310) | 2026-07-23 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1207 | Positive | Negative | Negative | Positive | 0.405 | -0.048 | -0.106 | 0.359 |  | 0 | Viability not assessed | 0 | 1.399 | 0.30 | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1208 | Negative | Negative | Negative | Negative | 5.9 | 4 | 3.8 | 5.1 | Median fluorescence intensity | 96.6 |  | 4.3 | 64.5 | >1.5X negative control median flourescence intensity | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1235 | Negative | Negative | Negative | Negative | ratio 1.30 | ratio 1.07 | ratio 1.13 | ratio 1.37 | MFI | 94 |  | MFI 7130 | MFI 30005 | ratio > 1.39 | 2026-07-21 | 2026-07-22 | Yes | 2026 |  |  | 4 | 2 | 2 | Yes
Lab-1256 | Not Tested | Negative | Negative | Not Tested |  | 0.278 | 0.275 |  | Median Log Channel | 85 |  | 0.309 | 5.66 | Ratio >1.8 | 2026-07-24 | 2026-07-27 | Yes | 2026 | Cell count too low | Without cells to repeat test | 2 | 2 | 2 | Yes
Lab-1305 | Positive | Negative | Negative | Negative | [10 empty] | 2026-07-23 | 2026-07-24 | Yes | 2026 | Technical failure | one result uninterpretable | 4 | 2 | 2 | Yes

## 2B-3-2026 B cells (G3)
Result summary | Serum 1 | Serum 2 | Serum 3 | Serum 4
Total distributed | 20 | 20 | 20 | 20
Total submitted | 20 | 20 | 20 | 20
Not Tested | 2 | 3 | 2 | 4
Positive | 18 | 12 | 0 | 12
Negative | 0 | 5 | 18 | 4
Equivocal | 0 | 0 | 0 | 0
% Positive | 100% | 70.59% | 0% | 75%
% Negative | 0% | 29.41% | 100% | 25%
% Equivocal | 0% | 0% | 0% | 0%
Consensus | Positive | No consensus | Negative | Positive
Unacceptable | Negative, Equivocal | N/A | Positive, Equivocal | Negative, Equivocal
Number acceptable | 18 | N/A | 18 | 12
Number unacceptable | 0 | N/A | 0 | 4
Assessed | Yes | No | Yes | Yes
Laboratory Code | B-cells Serum 1 | B-cells Serum 2 | B-cells Serum 3 | B-cells Serum 4 | B-cells Serum 1 Cytometer Reading | B-cells Serum 2 Cytometer Reading | B-cells Serum 3 Cytometer Reading | B-cells Serum 4 Cytometer Reading | Units for cytometer readings | Cell Viability | Comments | B-cells Local Negative Control | B-cells Local Positive Control | B-cells Positive Result Value | Date received | Date Tested | Submitted | Distribution Year | Reason for not testing | Confirm Reason | Tested | Correct | Assessed | Bcell
Lab-1038 | Positive | Positive | Negative | Positive | 5488 | 1699 | 235 | 1473 | fluorescence medians | 90 |  | 201 | 22327 | Ratio sample/negative control > 1.5 | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1041 | Not Tested | Not Tested | Not Tested | Not Tested | [6 empty] | No results, due the late delivery. |  |  |  | 2026-07-23 |  | Yes | 2026 | Sample delivery delay | No results, due late delivery. | 0 | 0 | 0 | Yes
Lab-1059 | Positive | Positive | Negative | Positive | 326.0 | 101.0 | 33.0 | 107.0 | Geo mean linear value | 85 |  | 40.0 | 443.0 | 3SD | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1068 | Positive | Positive | Negative | Positive | 11,9 | 3,33 | 1,23 | 2,63 | % shift |  |  | 1,0 | 5,6 | 70% increase in shift compared to negative control. | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1072 | Not Tested | Not Tested | Not Tested | Not Tested | [5 empty] | 95 |  |  |  |  | 2026-07-23 | 2026-07-24 | Yes | 2026 | Equivocal | equivoical | 0 | 0 | 0 | Yes
Lab-1074 | Positive | Not Tested | Negative | Positive | 3016 |  | 214 | 443 | median log channel | 85 |  | 176 | 25675 | 167 median log channel shift | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 3 | 3 | 3 | Yes
Lab-1100 | Positive | Negative | Negative | Negative | 1162 | 326 | 228 | 278 | median log channel | 99 |  | 258 | 11864 | NEG+3SD (358) | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 3 | Yes
Lab-1115 | Positive | Negative | Negative | Not Tested | [5 empty] | 89 |  |  |  |  | 2026-07-23 | 2026-07-23 | Yes | 2026 | Cell viability too low | cell viability too low | 3 | 2 | 2 | Yes
Lab-1127 | Positive | Positive | Negative | Positive | 9.2 | 4.4 | 1.3 | 3.1 | MFI | 97.7 |  | - | 50.6 | >2.5 | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1131 | Positive | Positive | Negative | Positive | 10659,1 | 3077,3 | 1367,3 | 2342,1 | Median Log Channel | 89 |  | 1271,4 | 97115,3 | 2SD | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1135 | Positive | Positive | Negative | Negative | 4889 | 909 | 322 | 746 |  | 95 |  | 460 | 78192 | 1,7 x higher than neg. control | 2026-07-23 | 2026-07-28 | Yes | 2026 |  |  | 4 | 2 | 3 | Yes
Lab-1152 | Positive | Positive | Negative | Positive | 5.120 | 1.120 | 0.586 | 1.010 |  |  |  | 0.566 | 2.410 | 3SD | 2026-07-27 | 2026-07-27 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1154 | Positive | Positive | Negative | Positive | 255709 | 76712 | 35239 | 66441 |  |  |  | 27336 | 148521 | A ratio of 2 as compared to the negative control replicates | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1169 | Positive | Negative | Negative | Not Tested | 3.07 | 0.903 | 0.508 |  | median log fluorescence intensity | 90 |  | 0.741 | 29.6 | 2SD | 2026-07-23 | 2026-07-27 | Yes | 2026 |  |  | 3 | 2 | 2 | Yes
Lab-1184 | Positive | Positive | Negative | Positive | 4587 | 1233 | 1090 | 4887 | median channel | 95 |  | 1190 | 9082 | >1.5 (1190) | 2026-07-23 | 2026-07-24 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1207 | Positive | Positive | Negative | Positive | 0.826 | 0.455 | -0.158 | 0.487 |  | 0 | Viability not assessed | 0 | 1.309 | 0.30 | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1208 | Positive | Positive | Negative | Positive | 287 | 69.9 | 20.6 | 52.7 | Median fluorescence intensity | 97.6 |  | 22.8 | 360.5 | >1.5X negative control Median fluorescence intensity | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1235 | Positive | Negative | Negative | Negative | ratio 6.84 | ratio 1.49 | ratio 1.01 | ratio 1.07 | MFI | 98 |  | MFI 13795 | MFI 104989 | ratio > 1.95 | 2026-07-21 | 2026-07-22 | Yes | 2026 |  |  | 4 | 2 | 3 | Yes
Lab-1256 | Positive | Positive | Negative | Positive | 7.96 | 1.62 | 0.201 | 1.82 | Median Log Channel | 90 |  | 0.414 | 14.9 | Ratio > 2.0 | 2026-07-24 | 2026-07-27 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1305 | Positive | Negative | Negative | Negative | [10 empty] | 2026-07-23 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 3 | Yes

## 2B-4-2026 B cells (G3)
Result summary | Serum 1 | Serum 2 | Serum 3 | Serum 4
Total distributed | 20 | 20 | 20 | 20
Total submitted | 20 | 20 | 20 | 20
Not Tested | 2 | 1 | 1 | 2
Positive | 10 | 0 | 0 | 17
Negative | 8 | 19 | 19 | 1
Equivocal | 0 | 0 | 0 | 0
% Positive | 55.56% | 0% | 0% | 94.44%
% Negative | 44.44% | 100% | 100% | 5.556%
% Equivocal | 0% | 0% | 0% | 0%
Consensus | No consensus | Negative | Negative | Positive
Unacceptable | N/A | Positive, Equivocal | Positive, Equivocal | Negative, Equivocal
Number acceptable | N/A | 19 | 19 | 17
Number unacceptable | N/A | 0 | 0 | 1
Assessed | No | Yes | Yes | Yes
Laboratory Code | B-cells Serum 1 | B-cells Serum 2 | B-cells Serum 3 | B-cells Serum 4 | B-cells Serum 1 Cytometer Reading | B-cells Serum 2 Cytometer Reading | B-cells Serum 3 Cytometer Reading | B-cells Serum 4 Cytometer Reading | Units for cytometer readings | Cell Viability | Comments | B-cells Local Negative Control | B-cells Local Positive Control | B-cells Positive Result Value | Date received | Date Tested | Submitted | Distribution Year | Reason for not testing | Confirm Reason | Tested | Correct | Assessed | Bcell
Lab-1038 | Positive | Negative | Negative | Positive | 1366 | 275 | 113 | 3664 | fluorescence medians | 90 |  | 200 | 16546 | Ratio sample/negative control > 1.5 | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1041 | Not Tested | Not Tested | Not Tested | Not Tested | [6 empty] | No results, due the late delivery. |  |  |  | 2026-07-23 |  | Yes | 2026 | Sample delivery delay | No results, due the late delivery. | 0 | 0 | 0 | Yes
Lab-1059 | Positive | Negative | Negative | Positive | 75.0 | 26.0 | 24.0 | 197.0 | Geo mean linear value | 80 |  | 26.0 | 406.0 | 3SD | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1068 | Positive | Negative | Negative | Positive | 2,07 | 1,5 | 0,76 | 8,22 | % shift |  |  | 1,0 | 5,38 | 70% increase in shift compared to negative control. | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1072 | Negative | Negative | Negative | Negative | 1577.25 | 1587.20 | 1426.70 | 1712.30 | MFI | 94 |  | 1316.90 | 13261.30 | 1000 log channel shift | 2026-07-23 | 2026-07-24 | Yes | 2026 |  |  | 4 | 2 | 3 | Yes
Lab-1074 | Not Tested | Negative | Negative | Positive |  | 193 | 187 | 1736 | median log channel | 85 |  | 203 | 12862 | 167 median log channel shift | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 3 | 3 | 3 | Yes
Lab-1100 | Negative | Negative | Negative | Positive | 270 | 229 | 222 | 813 | median log channel | 98 |  | 261 | 12246 | NEG+3SD (361) | 2026-07-24 | 2026-07-24 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1115 | Negative | Negative | Negative | Positive | [5 empty] | 92 |  |  |  |  | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1127 | Negative | Negative | Negative | Positive | 2.22 | 1.40 | 1.16 | 6.9 | MFI | 97.6 |  | - | 39.8 | >2.5 | 2026-07-23 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1131 | Positive | Negative | Negative | Positive | 1454,3 | 1037,8 | 665,2 | 6256,3 | Median Log Channel | 87 |  | 739,5 | 115081,2 | 2SD | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1135 | Positive | Negative | Negative | Positive | 631 | 285 | 250 | 1402 |  | 95 |  | 332 | 59182 | 1,7 x higher than neg. control | 2026-07-23 | 2026-07-28 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1152 | Negative | Negative | Negative | Positive | 0.782 | 0.671 | 0.581 | 1.610 |  |  |  | 0.636 | 1.260 | 3SD | 2026-07-27 | 2026-07-27 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1154 | Positive | Negative | Negative | Positive | 70296 | 26858 | 14141 | 156393 |  |  |  | 18097 | 36669 | A ratio of 2 as compared to the negative control replicates | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1169 | Negative | Negative | Negative | Positive | 0.503 | 0.405 | 0.395 | 2.66 | median log fluorescence intensity | 90 |  | 0.460 | 19.4 | 2SD | 2026-07-23 | 2026-07-27 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1184 | Positive | Negative | Negative | Positive | 1409 | 876 | 834 | 2169 | median channel | 95 |  | 1351 | 3033 | >1.5 (1351) | 2026-07-23 | 2026-07-24 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1207 | Positive | Negative | Negative | Positive | 0.566 | -0.084 | -0.282 | 0.767 |  | 0 | Viability not assessed | 0 | 1.352 | 0.30 | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1208 | Positive | Negative | Negative | Positive | 28.4 | 14.8 | 15.7 | 162.5 | Median fluorescence intensity | 96.6 |  | 14.9 | 226 | >1.5X negative control Median fluorescence intensity | 2026-07-22 | 2026-07-23 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1235 | Negative | Negative | Negative | Positive | ratio 1.03 | ratio 0.88 | ratio 0.72 | ratio 5.80 | MFI | 94 |  | MFI 11338 | MFI 55498 | ratio > 1.95 | 2026-07-21 | 2026-07-22 | Yes | 2026 |  |  | 4 | 3 | 3 | Yes
Lab-1256 | Positive | Negative | Negative | Positive | 0.796 | 0.336 | 0.226 | 4.52 | Median Log Channel | 85 |  | 0.311 | 7.54 | Ratio > 2.0 | 2026-07-24 | 2026-07-27 | Yes | 2026 | Cell count too low | Without cells to repeat test | 4 | 3 | 3 | Yes
Lab-1305 | Negative | Negative | Negative | Not Tested | [10 empty] | 2026-07-23 | 2026-07-24 | Yes | 2026 | Technical failure | one result uninterpretable | 3 | 2 | 2 | Yes