## 8-1-2026 Coeliac
8-1-2026 | Coeliac Disease
Distribution Date: 30/06/2026
Total distributed | 54
Total submitted | 46
Not tested | 1
Reference | DRB1*07:01, 07:01; DQA1*02:01, 02:01; DQB1*02:02, 03:03
Number acceptable | 45
Assessed | Yes
Comments
Lab Code | Coeliac Disease Results for Assessment | Coeliac Disease Alleles of Interest | Coeliac Disease Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-DQA1*02:01, -; HLA-DQB1*02:02, 03:03 | Specific genotype combinations of HLA-DQA1 and DQB1 alleles (HLA-DQA1*05, HLA-DQB1*02 / HLA-DQA1*03, HLA-DQB1*03:02 / HLA-DQA1*02:01, HLA-DQB1*02:02) | Positive for HLA-DQ2.2 The patient carries the HLA-DQ2.2 heterodimer (HLA-DQA1*02:01, HLA-DQB1*02:02) associated with Coeliac disease. Approximately 5% of HLA-DQ2.5 negative Coeliac disease sufferers carry this HLA type. The following considerations should be made when interpreting the HLA status of the patient: • Major histocompatibility antigens are expressed for life. Approximately 95% of Coeliac disease (CD) patients express the HLA-DQ heterodimers DQ2.5 (encoded by alleles DQB1*02 and DQA1*05) and DQ8 (encoded by alleles DQA1*03 and DQB1*03(DQ8)). • Expression of HLA-DQ2 and/or -DQ8 molecules is necessary, but not sufficient, to cause the disease. This result must be used in combination with further clinical information. • HLA-DQ2 and -DQ8 is present in approximately 40% of the UK population so a positive test in the absence of CD specific antibody testing has no predictive value, but a negative test can exclude a diagnosis of CD (NICE Clinical Guideline 86 2009). • Whilst HLA typing does not have an absolute diagnostic value it enables assessment of relative risk in the presence of a positive CD antibody test. References: Bodd et al. (2012) T-cell response to gluten in patients with HLA-DQ2.2 reveals requirement of peptide-MHC stability in celiac disease. Gastroenterology 142:552-561. D. Pritchard et al. (2024) UK NEQAS and BSHI guideline: Laboratory testing and clinical interpretation of HLA genotyping results supporting the diagnosis of Coeliac disease. International Journal of Immunogenetics 51:3-20. Medrano et al. (2012) HLA and celiac disease susceptibility: new genetic factors bring open questions about the HLA influence and gene-dosage effects. PLoS ONE 7(10):e48403. Megiorni et al. (2009) HLA-DQ status and risk gradient for celiac disease. Human Immunology 70:55-59. | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | HLA-DQA1*02 and HLA-DQB1*02 positive | DQA1*02, DQA1*03, DQA1*05 DQB1*02, DQB1*03:02 | The patient is HLA-DQ2.2 positive. Celiac disease is associated with this type of HLA that that occurs in 5% of the cases. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1008 | DQA1*02; DQB1*02:02/06/10/11/12++, *03:03/30/31/33/34++ | HLA-DQA1*05/DQB1*02 HLA-DQA1*03/DQB1*03:02 HLA-DQA1*02/DQB1*02 | This patient is negative for the HLA-DQ2.5 genotype (DQB1*02, DQA1*05), but positive for the HLA-DQ2.2 genotype (DQB1*02, DQA1*02), which has a weak association with Coeliac disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | DQA1*02 detected, DQA1*03 not detected, DQA1*05 not detected, DQB1*02 detected, DQB1*03:02 not detected | DQA1*02, DQA1*03, DQA1*05, DQB1*02, DQB1*03:02 | Subject carries the HLA-DQ2.2 allotype. This allotype is infrequently found in Coeliac disease (CD) patients, but an association cannot be excluded. However, the presence of this allotype should not be used to diagnose CD due to the very low positive predictive value of the association. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1020 | DQA1*02, DQB1*02 | DQ2.5 (DQA1*05, DQB1*02) DQ2.2 (DQA1*02, DQB1*02) DQ8 (DQA1*03, DQB1*03:02) |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1024 | HLA-DQA1*02 PRESENT HLA-DQA1*05 ABSENT HLA-DQB1*02 PRESENT HLA-DQB1*03:02(DQ8) ABSENT | HLA-DQA1*02,HLA-DQA1*05,HLA-DQB1*02 and HLA-DQB1*03:02(DQ8) | Présence du HLA-DQ2.2. Risque faible de prédisposition à la maladie cœliaque. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1025 | PREDISPOSED (HLA-DQ2.2/X) | HLA-DQA1*02 HLA-DQB1*02 (the kit also detected the HLA-DQB1-no*02 allele, indicating a heterozygous genotype at this locus) HLA-DRB1*07 | The HLA-DQ2.2/X genotype is associated with a low genetic risk for Coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | Yes
Lab-1026 | DQA1*05= negative; DQB1*02= positive; DQB1*03:02/05 (DQ8 group)= negative | DQA1*05; DQB1*02; DQB1*03:02/05 (DQ8 group) |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1027 | The sample contains DQ2.2 (HLA-DQA1*02; HLA-DQB1*02), genetic marker associated with celiac disease. |  |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | Yes
Lab-1031 | DQA1*05= negative; DQB1*02=positive; DQB1*03:02=negative | DQA1*05, DQB1*02, DQB1*03:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | Yes
Lab-1033 | HLA-DQB1*02: positive, HLA-DQA1*05:xx: negative, HLA-DQB1*03:02: negative | HLA-DQA1*05:xx, HLA-DQB1*02:01,02, HLA-DQB1*03:02 | The allel DQ2, which is weakly associated with Coeliac disease, could be detected. The allele DQ2.2 is not directly assessed, but cannot be excluded. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | Yes
Lab-1038 | HLA genotype : HLA-DQA1*02:01, DQA1*- (homozygous) HLA-DQB1*02:02, DQB1*03:03 Result : HLA-DQ2.5 (DQA1*05:01, DQB1*02:01, cis): absence HLA-DQ2.5 (DQA1*05:05, DQB1*02:02, trans) : absence HLA-DQ2.2 (DQA1*02:01, DQB1*02:02, cis) : presence HLA-DQ8 (DQA1*03, DQB1*03:02) : absence | HLA-DQ2.5 (DQA1*05:01, DQB1*02:01, cis) HLA-DQ2.5 (DQA1*05:05, DQB1*02:02, trans) HLA-DQ2.2 (DQA1*02:01, DQB1*02:02, cis) HLA-DQ8 (DQA1*03, DQB1*03:02) | This individual carries HLA-DQA1*02:01, DQB1*02:02 (DQ2.2) that has a weak association with Coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest Coeliac disease. >95% of CD patients carry HLA-DQ2 encoded by DQ2.5 (cis or trans) and 5% of CD patients carry HLA-DQ8 encoded by DQA1*03, DQB1*03:02. HLA genotyping has a good negative predictive value in the diagnosis of CD. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | Positive association with Coeliac disease: DQA1*02:01-DQB1*02:02 | DQA1 & DQB1 (DQ2/DQ8) | Presence of DQ2.2 | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | Yes
Lab-1057 | DQA1*02,DQB1*02,DQB1*03:03 | DQA1*05,DQB1*02,DQB1*03:02 | DQA1*05=absent DQB1*02=present DQB1*03:02=absent | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | HLA-DQA1*02:01 HLA-DQB1*02:XX/HLA-DQB1*03:XX (DQ9) |  |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | Yes
Lab-1065 | DQB1*02 heterozygous DQA1*05 negative DQB1*03:02 negative DQA1*03 negative | DQB1*02 (heterozygous/homozygous), DQA1*05, DQB1*03:02 and DQA1*03 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | DQ2.5: Negative DQ2.2: Positive DQ8: Negative | Combination of: DQ2.5: DQA1*05 and DQB1*02 DQ2.2: DQA1*02 and DQB1*02 DQ8: DQA1*03 and DQB1*0302 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1077 | DQA1*02:01:01,DQB1*02:02:01, DQB1*03:03:02 | DQA1 and DQB1 | DQB1*02, DQA1*05 (DQ2.5) Negative DQB1*02, DQA1*02 (DQ2.2) Positive DQB1*03:02 (DQ8) Negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | Yes
Lab-1079 | DQB1*02:02/02:06/02:10/02:11/02:12/02:20N/02:26/02:36/02:50/02:62/02:64/02:65/02:71/02:80/02:84/02:89/02:95/02:97/02:103/02:104/02:110/02:113/02:116/02:117/02:120/02:122/02:124/02:126/02:127/02:131/02:137/02:138/02:142/02:143/02:144/02:145/02:146/02:150/02:153/02:156/02:161/02:162N/02:165/02:167N/02:169/02:172/02:175/02:176N/02:179/02:183N/02:187/02:194N/02:195/02:199/02:200/02:201/02:203/02:204N/02:206N/02:207/02:210/02:211/02:212/02:217/02:221N/02:223/02:224/02:227/02:229/02:233/02:234/02:236/02:239/02:241/02:242/02:243/02:244/02:246/02:247/02:248/02:250/02:252/02:253/02:254/02:256/02:257/02:258/02:260 DQB1*03:03/03:03Q/03:30/03:31/03:33/03:34/03:39/03:41/03:43/03:79/03:86/03:87/03:88/03:89/03:91Q/03:96/03:97/03:98/03:99Q/03:104/03:105/03:111/03:117/03:123/03:124/03:126/03:136/03:145/03:149/03:168/03:176/03:200/03:209/03:212/03:222/03:227/03:230/03:248/03:249/03:258/03:270/03:280/03:304/03:313/03:316/03:319/03:332/03:357N/03:359/03:363/03:365/03:374/03:375N/03:382/03:384/03:393/03:395/03:398/03:402/03:405/03:406/03:414/03:445/03:453/03:461/03:477/03:478/03:487/03:489/03:494/03:505/03:507/03:509N/03:516/03:518/03:521/03:523/03:532/03:536/03:545/03:551N/03:568/03:574/03:575/03:583 DQA1*02:01/02:01Q/02:02N/02:03/02:04/02:05/02:07/02:08/02:09/02:10/02:11/02:12/02:13/02:14/02:15/02:16/02:18/02:19/02:20/02:21/02:22/02:23/02:24/02:25/02:27/02:28/02:29/02:30/02:31/02:32/02:33/02:35/02:36/02:37/02:38N/02:39/02:40/02:41/02:42/02:43/02:44/02:45/02:46/02:47/02:48/02:49/02:50/02:51 DQA1*02:01/02:01Q/02:02N/02:03/02:04/02:05/02:07/02:08/02:09/02:10/02:11/02:12/02:13/02:14/02:15/02:16/02:18/02:19/02:20/02:21/02:22/02:23/02:24/02:25/02:26/02:27/02:28/02:29/02:30/02:31/02:32/02:33/02:34/02:35/02:36/02:37/02:38N/02:39/02:40/02:41/02:42/02:43/02:44/02:45/02:46/02:47/02:48/02:49/02:50/02:51 | HLA-DQB1*02, HLA-DQB1*03:02 and HLA-DQA1*05 | This patient is POSITIVE for HLA-DQ2 (but is DQA1*05 NEGATIVE) and NEGATIVE for HLA-DQ8. Patients with this genotype have a LOW RISK of predisposition to Coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | DQB1*02,*03 DQA1*02 | DQA1, DQB1 | Presence of the beta chain of HLA dimer DQ2 (DQB1*02 positive, DQA1*05 negative) indicates low susceptibility for CD. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | Yes
Lab-1085 | NT | DQA1*02, DQA1*03, DQA1*05, DGA1*01/04/06, DQA1*02/03:01, DQA1*03:02/03, DQB1*02, DQB1*03:02, DQB1*03/06, DGB1*04, DQB1*05 | low yield, no result possible | Not Tested |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1086 | HLA DQ2.5: negative HLA DQ2.2: positive HLA DQ8: negative | DQA1*05, DQB1*02, DQA1*02, DQA1*03, DQB1*03, DQB1*0302 *0305, DQB1*0302 *0303 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | Yes
Lab-1090 | DQB1*02,*03(serological equivalent DQ9) | DQB1*02+DQA1*05 DQB1*02 DQB1*03:02 | Presence of a DQB1*02 allele not associated with the DQA1*05 : low to very low risk to develop Coeliac disease. This result alone does not confirm the diagnosis. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-22 | Yes
Lab-1104 | HLA Coeliac disease associated Heterodimer result: HLA-DQB1*02, DQA1*05 (DQ2.5) NOT detected HLA-DQB1*02, DQA1*02 (DQ2.2) DETECTED HLA-DQB1*03:02 (DQ8) NOT detected | HLA Coeliac Disease associated Heterodimers: HLA-DQB1*02, DQA1*05 (DQ2.5) HLA-DQB1*02, DQA1*02 (DQ2.2) HLA-DQB1*03:02 (DQ8) | This individual has a genotype which is associated with Coeliac disease (CD). More than 90% of individuals with Coeliac disease (CD), across all populations, express either HLA-DQ2 and / or DQ8 molecules with less than 2% lacking these human leukocyte antigen (HLA) types completely. The HLA-DQA1 and DQB1 alleles that encode these CD associated heterodimers (also referred to as DQ2.5, DQ2.2 and DQ8) are DQA1*05 DQB1*02, DQA1*02 DQB1*02 and DQB1*03:02. The majority of individuals with CD express the DQ2.5 heterodimer while remaining 5-10% express DQ2.2 or DQ8. As HLA-DQ2 and DQ8 are common in the general population (40% in Europeans) and only 1% develop CD, their presence has a low positive predictive value and does not confer a diagnosis of CD. A diagnosis should be based on serological testing from the gluten sensitivity profile, clinical findings and histological evaluation of duodenal biopsy. For further information see Pritchard, D. et al., 2024. UK NEQAS and BSHI guideline: Laboratory testing and clinical interpretation of HLA genotyping results supporting the diagnosis of Coeliac disease. International Journal of Immunogenetics, 51(Suppl. 1), 3–20. This assay has been performed by PCR-SSO (sequence specific oligonucleotides) based on Luminex xMAP technology using the LIFECODES HLA-DQA1/B1 SSO Typing kit (CE IVD) (Immucor), to determine the presence or absence of specific HLA alleles that form the above listed CD associated heterodimers. Heterodimer ambiguities have been resolved using additional PCR-SSP (sequence specific primers with the HLA Coeliac Disease Domino System (CE IVD) (Protrans), kit lot number 13-0 version 01 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-15 | Yes
Lab-1111 | HLA-DQ2.5 - negative (-) HLA-DQ2.2 - positive (+) HLA-DQ8 - negative (-) | HLA-DQ2.5 (DQA1*05/DQB1*02), HLA-DQ2.2 (DQA1*02/DQB1*02), HLA-DQ8 (DQA1*03/DQB1*03:02) | Haplotype: DQ2.2 Disease risk: low Only a negative result for the HLA-DQ2/DQ8 alleles rules out a genetic basis for celiac disease with 99% probability. A positive result does not confirm the disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1118 | DQB1*02:02, DQB1*03:03 DQA1*02:01, DQA1*02:01 | DQB1*02:01, DQB1*02:02, DQB1*03:02 and DQA1*05:01 | The patient possesses only one allele of the Coeliac Disease associated HLA-DQ2 molecule: DQB1*02 (DQA1*05 is not present). Patients with this genotype have a low risk of predisposition to Coeliac Disease. Other clinical indications are required for diagnosis. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1120 | • DQA1*05 Not Detected • DQB1*02 Detected • DQB1*0302 Not Detected | DQA1*05, DQB1*02, DQB1*0302 | This individual has one of the HLA-DQ variants associated with Coeliac disease (DQ2.2). More than 97% of Coeliac disease patients carry either HLA-DQ2.2, HLA-DQ2.5 or DQ8. However these variants are also present in approximately 40% of the general population and therefore whilst possession of the variant can support a diagnosis of Coeliac disease it is not per se diagnostic of the condition. This assay tests for the presence of HLA-DQ2.2 (DQB1*02), HLA-DQ2.5 (DQA1*05/DQB1*02) and HLA-DQ8 (DQB1*0302). | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | Yes
Lab-1132 | DQA1*05: negative; DQB1*02: positive; DQB1*03:02: negative. | DQA1*05 - DQB1*02 - DQB1*03:02 This test cannot detect variants other than those targeted by the analysis. Test sensitivity and specificity: 99%. | DQ2.2 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | Yes
Lab-1134 | HLA-DQB1*02 positive, HLA-DQA1*05 negative, HLA-DQB1*03:02 negative | HLA-DQB1*02, DQB1*03:02, HLA-DQA1*05 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1139 | DQA1*02:01, DQA1*02:01 DQB1*02:02, DQB1*03:03 | DQA1*05 + DQB1*02 (cis or trans) DQA1*XX + DQB1*03:02 DQA1*02 + DQB1*02 | This patient has DQA1*02, DQB1*02 (DQ2.2) that has a weak association with Coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest Coeliac disease. The presence of an associated HLA genotype does not confer a diagnosis of Coeliac disease and has a low positive predictive value for Coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | DQB1*03:02: negative DQA1*05: negative DQB1*02: positive DQA1*02: positive. Reported serotypes: DQ2.2 | HLA-DQB1*03:02, HLA-DQA1*05, HLA-DQB1*02, HLA-DQA1*02 | "HLA-DQ2.2 is detected in the form of HLA-DQB1*02 and HLA-DQA1*02. A small minority of Coeliac patients has these alleles. The alleles are common in the general population. Coeliac disease is unlikely. However, the test alone cannot exclude Coeliac disease." | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | Yes
Lab-1146 | DQA1, DQA1 DQB1*02, DQB1 | DQA1*05 DQB1*02, DQB1*03:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1150 | DQ2.2 | DQA1*05-DQB1*02-DRB1*03 DQA1*05-DQB1*03:01-DRB1*11/DRB1*12 DQA1*02-DQB1*02-DRB1*07 DQA1*03-DQB1*03:02-DRB1*04 | detected HLA genotype is associated with a rare risk of Coeliac disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | Yes
Lab-1151 | DQA1*02:01, 02:01 DQB1*02:02, 03:03 | DQA1*05 , DQA1*02, DQB1*02, DQB1*03:02 | This individual carries HLA-DQA1*02, DQB1*02:02 (DQ2.2) that has a weak association with Coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest Coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | Yes
Lab-1154 | DQ2 present | DQ2, DQ8 and DQ7 and DQA05 | DQ2 present that is associated with Coeliac disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1176 | HLA-DQA1*02:01, 02:01; HLA-DQB1*02:02, 03:03 | HLA-DQA1*05:01/DQB1*02:01 (DQ2.5); HLA-DQA1*03/DQB1*03:02 (DQ8); HLA-DQA1*02:01/DQB1*02:02 (DQ2.2) | Positive for DQ2.2. This individual carries HLA-DQA1*02, DQB1*02:02 (DQ2.2) that has a weak association with Coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest Coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1186 | DQA1*05 - NEGATIVE DQB1*02 - POSITIVE DQA1*03 - NEGATIVE DQB1*03:02 - NEGATIVE DQB1*03:03 - POSITIVE | HLA-DQA1*05, HLA-DQA1*03, HLA-DQB1*02, HLA-DQB1*03:02, HLA-DQB1*03:03 | Low risk of celiac disease | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1199 | DQB1*02-DQA1*02 DQB1*03:03-DQA1*02 | HLA-DQ2.5/DQ8 | TThe presence of HLA-DQA1*02/DQB1*02:02 (DQ2.2) has a weak association with Coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest Coeliac disease. The presence of an associated HLA genotype does not confer a diagnosis of Coeliac disease and has a low positive predictive value for Coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | HLA DQ2.2 = Present ; HLA DQ2.5 = Absent ; HLA DQ8 = Absent | DQA1*02; DQA1*02/*0301; DQA1*03; DQA1*0302/03; DQA1*05; DQB1*02; DQB1*02/*0302. |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | Genotypen er associeret med cøliaki. Analysen har påvist HLA-DQB1*02/HLA-DQA1*02. Den genetiske risiko for cøliaki er til stede. | DQA1*05 negative DQB1*02 positive DQA1*02 positive DQB1*03:02 negative DQA1*03 negative |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | Yes
Lab-1208 | DQB1*02 - Positive DQA1*05 - Negative DQB1*03:02 - Negative | DQB1*02 DQA1*05 DQB1*03:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | Yes
Lab-1209 | HLA-DQ2.2 - POSITIVE HLA-DQ2.5 - NEGATIVE HLA-DQ8 - NEGATIVE b-subunit HLA-DQ2.2/DQ2.5 - POSITIVE |  |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | Yes
Lab-1215 | DQB1*03:02 - Negative DQB1*02 - Positive DQA1*05 - Negative | DQB1*03:02 DQB1*02 DQA1*05 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1239 | This patient is POSITIVE for the HLA-DQ2.2 genotype but negative for HLA-DQ2.5 and -DQ8. | DQA1*02, *03, *05 and DQB1*02, *03:02 | CD is associated with HLA heterodimers HLA-DQA1*05-DQB1*02 (DQ2.5), HLA-DQA1*03-DQB1*03:02 (DQ8) or HLA-DQA1*02-DQB1*02 (DQ2.2). The absence of any of these associated genotypes excludes the diagnosis of Coeliac disease with a high probability (<1% likelihood). Further information is available in the ESsCD guidelines (2019): https://doi.org/10.1177/2050640619844125. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1265 | DQB1*02:02 ; DQA1*02:01 DQB1*03:03 ; DQA1*02:01 | DQB1*02:01/DQA1*05:01 ; DQB1*03:02/DQA1*03 DQB1*02:02/DQA1*02:01 | Investigation of susceptibility to celiac disease PRESENCE of the HLA-DQB1*02:02/HLA-DQA1*02:01(DQ2.2) haplotype associated with a lower risk of developing celiac disease. This result indicates celiac disease only if it is associated with clinical, biological and radiological signs corresponding to the diagnostic criteria. Absence of the DQ8 haplotype (HLA-DQB1*03:02/HLA-DQA1*03) (Réf. : Ludvig M.Sollid et al., 2012 ; Megiorni F et al. 2012 ; Hadithi M. et al. 2007 ; Margaritte-Jeannin P et al., 2004) | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1268 |  |  |  | Unacceptable | No results returned | No | [5 empty] | Yes
Lab-1269 | Presence of DQB1*02, absence of DQA1*05, DQB1*03:02, DQA1*03 | DQB1*02, DQA1*05, DQB1*03:02, DQA1*03 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | Yes
Lab-1272 | HLA-DQA1*05 -> negative HLA-DQB1*02 -> positive HLA-DQB1*03:02 -> negative | HLA-DQA1*05, HLA-DQB1*02, HLA-DQB1*03:02 | DQ2.2 +; DQ8 - Geringes Risiko einer Glutenunverträglichkeit | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | Yes
Lab-1284 | DQB1*02 – DQA1*02 | HLA-DQB1*03:02; HLA-DQB1*02 – HLA-DQA1*05; HLA-DQB1*02 – HLA-DQA1*02 | Positive for HLA-DQ2.2, which is associated with Coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | Yes
Lab-1291 | DRB1*07 - DQA1*02:01 - DQB1*02 DRB1*07 - DQA1*02:01 - DQB1*03:03 |  | DR7 DQ2 DR7 DQ9 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-10 | Yes
Lab-1295 | DQ2.2 | // | // | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | Yes
Lab-1296 | DQB1*02 DQA1*02 DRB1*07 DQB1*0302 | DQB1*02 DQA1*02 (DRB1*07, DQB1*0302) | HLA-DQ2.2 | Unacceptable | Reported DQB1*03:02 rather than DQB1*03:03 | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1304 | HLA-DQA1*05 Negative HLA-DQB1*02 Positive/Heterozygous HLA-DQB1*03:02 Negative | HLA-DQA1*05 HLA-DQB1*02 HLA-DQB1*03:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | Yes
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-1-2026 Narcolepsy
8-1-2026 | Narcolepsy
Distribution Date: 30/06/2026
Total distributed | 25
Total submitted | 52
Not tested | 1
Reference | DQB1*02:02, 03:03
Number acceptable | 24
Assessed | Yes
Comments
Lab Code | Narcolepsy Results for Assessment | Narcolepsy Alleles of Interest | Narcolepsy Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-DQB1*02:02, 03:03 | HLA-DQB1*06:02 | This patient is Negative for HLA-DQB1*06:02, the allele associated with Narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | DQB1*06:02 negative | DQB1*06:02 | The patient is lacking the HLA-type (DQB1*06:02) wich is associatetd with narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1008 | DQB1*02:02/06/10/11/12++, *03:03/30/31/33/34++ | HLA-DQB1*06:02 | This patient does not have the HLA-DQB1*06:02 allele associated with narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | We didn't find narcolepsy associated alleles in sample 01. | DQA1*01:02, DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1024 | DQB1*06:02 ABSENT | DQB1*06:02 | Absence de l'allèle DQB1*06:02. Le diagnostic de narcolepsie de type 1 est peu probable mais n’est pas formellement exclu. L’interprétation de ce résultat doit être réalisée en fonction du contexte clinique. Références : - Rennac V et al. Guidelines from the French-speaking Society for Histocompatibility and Immunogenetics (SFHI) for harmonisation of HLA genotyping in autoimmune diseases, drug hypersensitivity and pharmacogenetics. HLA 2025;106:e70442. - Capittini C et al. Correlation between HLA-DQB1*06 :02 and narcolepsy with and without cataplexy. Sleep Medicine 2018;52:150-157. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | DQB1*06:02= negative | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | No tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | HLA-DQB1*06:02: negative | HLA-DQB1*06:02 | In absence of the Allele HLA-DQB1*06:02, narcolepsy type I with cataplexy can be excluded with a sensitivity of approximately 90%. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | Yes
Lab-1038 | HLA genotype : HLA-DQB1*02:02, DQB1*03:03 Result: HLA-DQB1*06:02 : absence | HLA-DQB1*06:02 | This individual has no DQB1* allele associated with narcolepsy. The HLA-DQB1*06:02 allele is found in 15-25% of the overall population and in 90-100% of narcolepsy patients. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | No known associatio with narcolepsy | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | Yes
Lab-1057 | DQB1*02,DQB1*03 | DQB1*06:02 | DQB1*06:02=absent | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | DQB1*06:02 negative | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | DQB1*02:02:01, DQB1*03:03:02 | DQB1*06:02 | DQB1*06:02 Negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | Yes
Lab-1079 | DQB1*02:02/02:06/02:10/02:11/02:12/02:20N/02:26/02:36/02:50/02:62/02:64/02:65/02:71/02:80/02:84/02:89/02:95/02:97/02:103/02:104/02:110/02:113/02:116/02:117/02:120/02:122/02:124/02:126/02:127/02:131/02:137/02:138/02:142/02:143/02:144/02:145/02:146/02:150/02:153/02:156/02:161/02:162N/02:165/02:167N/02:169/02:172/02:175/02:176N/02:179/02:183N/02:187/02:194N/02:195/02:199/02:200/02:201/02:203/02:204N/02:206N/02:207/02:210/02:211/02:212/02:217/02:221N/02:223/02:224/02:227/02:229/02:233/02:234/02:236/02:239/02:241/02:242/02:243/02:244/02:246/02:247/02:248/02:250/02:252/02:253/02:254/02:256/02:257/02:258/02:260 DQB1*03:03/03:03Q/03:30/03:31/03:33/03:34/03:39/03:41/03:43/03:79/03:86/03:87/03:88/03:89/03:91Q/03:96/03:97/03:98/03:99Q/03:104/03:105/03:111/03:117/03:123/03:124/03:126/03:136/03:145/03:149/03:168/03:176/03:200/03:209/03:212/03:222/03:227/03:230/03:248/03:249/03:258/03:270/03:280/03:304/03:313/03:316/03:319/03:332/03:357N/03:359/03:363/03:365/03:374/03:375N/03:382/03:384/03:393/03:395/03:398/03:402/03:405/03:406/03:414/03:445/03:453/03:461/03:477/03:478/03:487/03:489/03:494/03:505/03:507/03:509N/03:516/03:518/03:521/03:523/03:532/03:536/03:545/03:551N/03:568/03:574/03:575/03:583 | HLA-DQB1*06:02 | DQB1*06:02 NEGATIVE. Narcolepsy is associated with the expression of the human leukocyte antigen (HLA) class II molecule DQB1*06:02. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | DQB1*02,*03(serological equivalent DQ9) | DQB1*06:02 | Absence of the susceptibility allele for narcolepsy-cataplexy DQB1*06:02. This allele is present in 85 to 95% of the patients with narcolepsy-cataplexy, in 40 to 60% of patients with narcolepsy without cataplexy and in 18% of patients with idiopathic hypersomnia but it is also present in 12 to 38% of the population. This result makes the diagnosis of narcolepsy-cataplexy unlikely but does not exclude the diagnosis. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-22 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-15 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | DQB1*02:02, DQB1*03:03 | DQB1*06:02 | HLA-DQB1*06:02 is associated with narcolepsy-cataplexy. This patient is NEGATIVE for HLA-DQB1*06:02. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | HLA-DQB1*06:02 negative | HLA-DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1139 | DQB1*02:02, DQB1*03:03 | DQB1*06:02 | This patient does not have the HLADQB1*06 allele classically associated with Narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-DQB1*02:02, 03:03 | HLA-DQB1*06:02 | HLA-DQB1*06:02 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | DQB1*02-DQA1*02 DQB1*03:03-DQA1*02 | HLA-DQB1*06:02 | Patient has tested negative for HLA DQB1*06:02 allele which is associated with Narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | DQB1*06:02 - Negative | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | Yes
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | HLA-DQB1*06:02 NEGATIVE | DQB1*06:02 | The HLA-DQB1*06:02 allele associated with, but not diagnostic of, Narcolepsy is absent in this patient. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1265 | DQB1*02:02 DQB1*03:03 | DQB1*06:02 | Investigation of Susceptibility to Narcolepsy ABSENCE of the HLA-DQB1*06:02 susceptibility allele described in narcolepsy-cataplexy. This result makes the diagnosis of narcolepsy-cataplexy unlikely, but does not exclude the diagnosis. It should be compared with the clinical and biological signs and the functional tests corresponding to the diagnostic criteria. (Réf : Correlation between HLA-DQB1*06:02 and narcolepsy with and without cataplexy: approving a safe and sensitive genetic test in four major ethnic groups. A Systematic meta analysis, Sleep Medecine, 2018) | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | Absence of DQB1*06:02 | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | Yes
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | DQB1*06:02 negative. | HLA-DQB1*06:02 | Genotype not associated with narcolepsy. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | Yes
Lab-1291 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-10 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-1-2026 Actinic Prurigo
8-1-2026 | Actinic Prurigo
Distribution Date: 30/06/2026
Total distributed | 6
Total submitted | 53
Not tested | 0
Reference | DRB1*07:01, 07:01
Number acceptable | 6
Assessed | Yes
Comments
Lab Code | Actinic Prurigo Results for Assessment | Actinic Prurigo Alleles of Interest | Actinic Prurigo Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-DRB1*07:01, -. | HLA-DRB1*04:07 | This patient is Negative for HLA-DRB1*04:07, the allele associated with Actinic Prurigo | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | DRB1*07:01/03/05/06/07++ | HLA-DRB1*04:07 | This patient does not have the HLA-DRB1*04:07 allele associated with Actinic Prurigo | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | No tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | No
Lab-1038 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1067 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | No
Lab-1079 | DRB1*07:01/07:03/07:05/07:06/07:07/07:08/07:09/07:10N/07:12/07:13/07:14/07:15/07:16/07:17/07:18/07:19/07:21/07:22/07:23/07:24/07:25/07:27/07:28/07:29/07:30/07:31/07:32/07:33/07:34/07:35/07:36/07:37/07:38/07:39/07:40/07:41/07:42/07:43/07:44/07:45/07:46/07:47/07:48/07:49/07:50/07:51/07:52/07:53/07:54/07:55/07:56/07:57/07:58N/07:59/07:60/07:61/07:62/07:63/07:64/07:65/07:66/07:67/07:68N/07:69/07:70/07:71/07:72/07:73/07:74/07:75/07:76/07:77/07:78/07:79/07:80/07:81/07:82/07:84/07:85/07:86/07:87N/07:88/07:89/07:90/07:91/07:92/07:93/07:94/07:95/07:96/07:97/07:98/07:99/07:100/07:102/07:103/07:104/07:105/07:106/07:107/07:108/07:109/07:110/07:111/07:112/07:113/07:114/07:115/07:116/07:117/07:118N/07:119/07:120/07:121/07:122/07:123/07:124/07:125/07:126/07:127/07:128/07:129N/07:130/07:131/07:132/07:133/07:134/07:135/07:136/07:137/07:138/07:139/07:140/07:141/07:142/07:143N/07:144/07:145/07:146/07:147N/07:148/07:149N/07:150/07:151/07:152/07:153/07:154/07:155/07:156N/07:157/07:158/07:159/07:160/07:161/07:162/07:163N/07:164/07:165/07:166/07:167/07:168/07:169/07:170/07:171/07:172 DRB1*07:01/07:03/07:05/07:06/07:07/07:08/07:09/07:10N/07:12/07:13/07:14/07:15/07:16/07:17/07:18/07:19/07:21/07:22/07:23/07:24/07:25/07:26N/07:27/07:28/07:29/07:30/07:31/07:32/07:33/07:34/07:35/07:36/07:37/07:38/07:39/07:40/07:41/07:42/07:43/07:44/07:45/07:46/07:47/07:48/07:49/07:50/07:51/07:52/07:53/07:54/07:55/07:56/07:57/07:58N/07:59/07:60/07:61/07:62/07:63/07:64/07:65/07:66/07:67/07:68N/07:69/07:70/07:71/07:72/07:73/07:74/07:75/07:76/07:77/07:78/07:79/07:80/07:81/07:82/07:84/07:85/07:86/07:87N/07:88/07:89/07:90/07:91/07:92/07:93/07:94/07:95/07:96/07:97/07:98/07:99/07:100/07:101N/07:102/07:103/07:104/07:105/07:106/07:107/07:108/07:109/07:110/07:111/07:112/07:113/07:114/07:115/07:116/07:117/07:118N/07:119/07:120/07:121/07:122/07:123/07:124/07:125/07:126/07:127/07:128/07:129N/07:130/07:131/07:132/07:133/07:134/07:135/07:136/07:137/07:138/07:139/07:140/07:141/07:142/07:143N/07:144/07:145/07:146/07:147N/07:148/07:149N/07:150/07:151/07:152/07:153/07:154/07:155/07:156N/07:157/07:158/07:159/07:160/07:161/07:162/07:163N/07:164/07:165/07:166/07:167/07:168/07:169/07:170/07:171/07:172 | HLA-DRB1*04:07 | HLA-DRB1*04:07 NEGATIVE. Actinic Prurigo is associated with the expression of the human leukocyte antigen (HLA) class II molecule DRB1*04:07 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-22 | No
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-15 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | DRB1*07:01, DRB1*07:01 | DRB1*04:07 | This patient does not have the DRB1*04:07 allele associated with Actinic Prurigo | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-DRB1*07:01, 07:01 | HLA-DRB1*04:07 | HLA-DRB1*04:07 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | DRB1*07 DRB1*07 | HLA-DRB1*04:07 | Negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1265 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | No
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-10 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | [5 empty] | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | No

## 8-1-2026 Birdshot
8-1-2026 | Birdshot
Distribution Date: 30/06/2026
Total distributed | 16
Total submitted | 52
Not tested | 1
Reference | A*01:01, 29:102
Number acceptable | 15
Assessed | Yes
Comments
Lab Code | Birdshot Results for Assessment | Birdshot Alleles of Interest | Birdshot Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-A*01:01, 29:102 | HLA-A29 | This patient is Positive for HLA-A29, the specificity associated with Birdshot Retinopathy | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | A*01:01/03/04/09/11N++, *29:02/06/08N/09/10++ | HLA-A*29 | This patient appears to have the HLA-A29 antigen associated with birdshot retinopathy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | HLA-A*29: positive. | HLA-A*29 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | Yes
Lab-1038 | HLA Genotype: HLA-A*01:01, A*29:102 Result: HLA-A*29 : presence of a rare allele of HLA-A*29 which was not studied for its related risk of Birdshot chorioretinitis | HLA-A*29 | The individual carries a rare allele of HLA-A*29 which was not studied for its related risk of Birdshot chorioretinitis. The presence of the HLA-A*29 allele (A*29:01 or A*29:02), associated with the characteristic clinical signs of the disease, would strongly support the diagnosis of Birdshot chorioretinitis. The prevalence of HLA-A29 in patients with Birdshot chorioretinitis is 90–100% according to published studies. No specific preventive measures, nor family or prenatal screening, are currently recommended for carriers of this susceptibility allele. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | A*01,A*29 | A*29 | A*29=present | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | A*29 positive | A*29 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | A*01:01:01, A*29:102 | A*29 | A*29 Positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | Yes
Lab-1079 | A*01:01/01:01L/01:01N/01:03/01:04N/01:09/01:11N/01:17/01:18N/01:22N/01:23/01:24/01:29/01:32/01:33/01:35/01:36/01:37/01:38/01:39/01:40/01:42/01:44/01:45/01:47/01:48/01:50/01:52N/01:53N/01:54/01:55/01:56N/01:57N/01:58/01:59/01:61/01:62/01:65/01:67/01:70/01:74/01:75/01:76/01:77/01:78/01:79/01:80/01:81/01:82/01:84/01:86/01:87N/01:88/01:90/01:91/01:92/01:93/01:94/01:96/01:97/01:100/01:101/01:103/01:104/01:105/01:106/01:107/01:108/01:109/01:110/01:111/01:112/01:113/01:115/01:116/01:117/01:118/01:120/01:121/01:122/01:123N/01:124/01:125/01:129/01:131/01:132/01:135/01:137/01:138/01:139/01:140/01:141/01:142/01:143/01:144/01:145/01:146/01:148/01:150/01:151/01:152/01:153/01:154/01:155/01:156/01:157/01:158/01:161/01:162N/01:164/01:165/01:166/01:170/01:171/01:172/01:174/01:175/01:177/01:179N/01:180/01:181/01:182/01:183/01:184/01:185/01:186N/01:187/01:189/01:191/01:193/01:195/01:197/01:198/01:199/01:201/01:202/01:203/01:204/01:206/01:207/01:210/01:211/01:212/01:213/01:216/01:217/01:218/01:219/01:220/01:221/01:222/01:223/01:224/01:225/01:226/01:227/01:228Q/01:230/01:231/01:232/01:234/01:235/01:237/01:239/01:240N/01:241/01:242/01:245/01:246/01:247N/01:248Q/01:249/01:251/01:252/01:253/01:254/01:257/01:258N/01:259/01:260/01:261/01:262/01:263/01:265/01:266/01:268/01:269N/01:270/01:271/01:272/01:273/01:274/01:275/01:276/01:277/01:278/01:280/01:281Q/01:282/01:283/01:285N/01:287N/01:288/01:291/01:292/01:293N/01:294/01:295/01:296/01:297/01:300/01:303/01:305/01:306/01:308N/01:309/01:311/01:312/01:313/01:315/01:316/01:317/01:319/01:320N/01:321/01:322/01:323/01:324/01:325/01:326N/01:327/01:328N/01:329/01:330/01:331/01:332/01:333/01:334/01:335/01:336N/01:337/01:338/01:339/01:341/01:342/01:343/01:345/01:346/01:347/01:348/01:349/01:350/01:351/01:352/01:353/01:354/01:356/01:357/01:358/01:359/01:361N/01:362/01:365/01:366/01:367/01:368/01:369/01:370/01:371/01:372/01:373/01:374/01:375/01:376/01:377/01:378/01:379N/01:382/01:383/01:384N/01:385/01:386/01:387/01:388/01:389/01:392/01:393/01:394N/01:395/01:400N/01:401/01:402/01:404/01:406/01:407/01:408/01:409/01:410/01:411N/01:414/01:415/01:416/01:417N/01:419/01:421N/01:422/01:423/01:424/01:425/01:426/01:428/01:429/01:430/01:431/01:432/01:433N/01:434/01:435/01:436Q/01:437/01:438/01:439/01:440/01:441/01:442/01:443/01:444/01:445/01:446/01:447/01:448/01:450N/01:451N/01:452N/01:456/01:457/01:458/01:459/01:460/01:461/01:463/01:465/01:467/01:468/01:469/01:470 A*29:02/29:06/29:08N/29:09/29:10/29:11/29:21/29:23/29:25/29:26/29:27/29:29/29:30/29:32/29:34/29:36/29:42/29:43/29:44/29:45/29:46/29:47/29:50/29:52/29:53/29:54/29:59/29:63/29:65/29:66/29:68/29:69/29:70/29:72/29:74/29:75/29:78N/29:85/29:88/29:91/29:92/29:93/29:94/29:95/29:96/29:100/29:102/29:103/29:106/29:108/29:109/29:111/29:112N/29:114/29:116/29:119/29:120/29:121/29:122/29:126Q/29:127/29:130/29:131/29:133/29:134/29:135/29:136/29:137/29:138/29:139/29:140/29:141/29:143/29:145/29:146/29:148/29:149/29:156/29:160/29:161/29:165/29:167/29:168/29:173/29:175Q/29:177/29:178/29:179/29:180/29:181/29:184/29:186/29:188/29:189/29:191/29:193/29:194/29:197/29:198 | A*29 | HLA-A*29 POSITIVE. Birdshot retinochoroidopathy is associated with the expression of the human leukocyte antigen (HLA) class I molecule A*29. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | NT | HLA-A29 | low yield, no result possible | Not Tested |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | A*01,*29 | A*29 | Presence of the A*29 susceptibility allele. The presence of the A*29 allele, associated with the clinical signs of the disease, is strongly in favor of the diagnosis of birdshot disease. The prevalence of A*29 in patients with birdshot is 90 to 100% according to published studies. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-22 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-15 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | A*01:01, A*29:102 | A*29 | The patient is positive for HLA-A29 which is associated with Birdshot chorioretinopathy. The presence of HLA-A29 alone cannot be used to diagnose disease. It must be used with other clinical indications. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | A*01, A*29 | A*29 | This patient has a HLA genotype associated with Birdshot Retinopathy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-A*01, 29 | HLA-A*29 | HLA-A*29 positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | A*01 A*29 | HLA-A*29 | Positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | HLA-A29 POSITIVE | A*29 | The presence of the HLA-A29 genotype is associated with, but not diagnostic of, Birdshot Chorioretinopathy. HLA-A29 is present in ~9% of the UK population. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1265 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | Presence of A*29 | A*29 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | Yes
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-10 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | [5 empty] | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | No

## 8-1-2026 Behcets
8-1-2026 | Behcets
Distribution Date: 30/06/2026
Total distributed | 20
Total submitted | 51
Not tested | 2
Reference | B*44:03, 57:01
Number acceptable | 18
Assessed | Yes
Comments
Lab Code | Behcets Results for Assessment | Behcets Alleles of Interest | Behcets Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-B*44:03, 57:01 | HLA-B51 | This patient is Negative for HLA-B51, the specificity associated with Behcet's Disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | B*44:03/26/30/35/36++, *57:01/18/21/23/27++ | HLA-B*51 | This patient does not have the HLA-B51 antigen associated with Behcet’s disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | HLA-B*51 ABSENT | HLA-B*51 | « Ce résultat n'exclut pas le diagnostic de maladie de Behçet : 30 à 50% des patients atteints ne possèdent pas l’allèle HLA B*51 ». | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | HLA-B*51: negative. | HLA-B*51 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | Yes
Lab-1038 | HLA genotype : HLA-B*44:03, B*57:01 Result : HLA-B*51: absence | HLA-B*51 | Absence of the HLA-B*51 susceptibility allele This result does not rule out the diagnosis of anterior inflammatory uveitis associated with Behçet's disease. (Ref.: Zeidan MJ et al., Behçet's disease pathophysiology: a contemporary review. Auto Immun Highlights. 2016) | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | B*44,B*57 | B*51 | B*51=absent | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | B*51 negative | B*51 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | B*44:03:01, B*57:01:01 | B*51 | B*51 Negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | Yes
Lab-1079 | B*44:03/44:26/44:30/44:35/44:36/44:37/44:38/44:39/44:57/44:60/44:85/44:89/44:92/44:94/44:98/44:103/44:108N/44:111/44:114/44:115/44:120/44:122/44:125/44:128/44:141/44:147/44:154/44:155/44:157/44:159/44:160Q/44:161/44:164/44:165/44:167/44:174/44:175/44:178/44:180/44:182/44:183/44:186/44:188/44:189/44:192/44:194/44:198N/44:202/44:204/44:205/44:207/44:215/44:222/44:228/44:231/44:233/44:234/44:237N/44:239/44:245/44:248/44:250/44:252/44:258/44:266/44:278/44:280/44:281/44:284/44:286/44:295/44:298/44:300/44:302/44:304/44:309N/44:310N/44:318/44:319/44:323/44:324/44:325/44:328N/44:336/44:346/44:348/44:349/44:354N/44:367/44:373/44:378/44:379/44:381/44:389/44:390/44:397/44:398/44:403/44:408/44:409/44:411/44:412/44:419/44:425/44:428/44:429N/44:433/44:443/44:450/44:456/44:460/44:466N/44:471/44:472/44:474/44:476/44:483/44:485/44:488/44:491/44:493/44:500/44:501/44:507/44:508/44:509/44:510/44:511/44:512N/44:513/44:514/44:516/44:522/44:524/44:527/44:530/44:533/44:536/44:542/44:546/44:547/44:548/44:551/44:552 B*57:01/57:18/57:21/57:23/57:27/57:29/57:31/57:33/57:35/57:37/57:38/57:40/57:43/57:44/57:45/57:47/57:48/57:50/57:51/57:52/57:54/57:55/57:58/57:61/57:64/57:65/57:67/57:68/57:69/57:71/57:72/57:73/57:74/57:75/57:77/57:79N/57:85/57:86/57:87/57:88/57:90/57:91/57:92/57:93/57:95/57:97/57:99/57:100/57:102/57:103/57:105/57:106/57:108/57:110/57:111/57:113/57:114/57:115/57:116/57:117/57:121/57:122N/57:123/57:125/57:126/57:127/57:129/57:130N/57:132/57:133/57:134/57:135/57:136/57:138/57:139N/57:141/57:142N/57:143N/57:144/57:145/57:146/57:149/57:150/57:151/57:152/57:153/57:154/57:156/57:157/57:158/57:159/57:162N/57:163/57:164/57:165/57:166/57:167/57:168/57:169/57:171/57:172/57:173/57:176/57:177/57:178/57:179/57:181N/57:182/57:183/57:184/57:187/57:188/57:190/57:191/57:192/57:194/57:195/57:197/57:198/57:199 | HLA-B*27 and HLA-B*51 | HLA-B*27 NEGATIVE, HLA-B*51 NEGATIVE Behcet's disease is associated with the expression of the human leukocyte antigen (HLA) class I molecule B51. HLA-B27 result can be used for differential diagnoses due to the overlapping clinical features of Behcets and HLA-B27 related diseases. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | NT | HLA-B51 | low yield, no result possible | Not Tested |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | B*44,*57 | B*51 | Absence of the susceptibility allele HLA-B*51. This result does not exclude the diagnosis of Behçet's disease: 30-50% of patients do not have the B*51 allele. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-22 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-15 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | B*44:03, *57:01 | B*51 | The patients HLA type does NOT include HLA - B51 which is associated with Behcet’s disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | B*44, B*57 | B*51 | This patient does not have a HLA genotype associated with Behcet's Disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-B*44:03, 57:01 | HLA-B*51 | HLA-B*51 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1186 | B*51 - NEGATIVE | HLA-B*51 | Low risk of Behçet Syndrome | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1199 | B*44 B*57 | HLA-B*51 | Negative The patient does not have the HLA antigen (HLA-B51) associated with Behcets | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | HLA-B*51 - Negative | HLA-B*51 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1239 | HLA-B51 NEGATIVE | B*51 | The HLA-B51 genotype associated with, but not diagnostic of Behcet's disease is absent in this patient. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1265 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | Absence of B*51 | B*51 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | Yes
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-10 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-1-2026 Rheumatoid
8-1-2026 | Rheumatoid
Distribution Date: 30/06/2026
Total distributed | 6
Total submitted | 52
Not tested | 1
Reference | DRB1*07:01, 07:01
Number acceptable | 5
Assessed | Yes
Comments
Lab Code | Rheumatoid Results for Assessment | Rheumatoid Alleles of Interest | Rheumatoid Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | No
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | No
Lab-1038 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | DRB1*07 | DRB1*04 | DRB1*04=absent | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1067 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | No
Lab-1079 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | DRB1*07:01,- | DRB1*04:01,*04:04,*04:05,*04:08 DRB1*10 DRB1*01:01,*01:02,*01:04 DRB1*14:06 | Absence of any susceptibility epitope carried by an HLA-DRB1 molecule. The patient has no epitope associated with a risk of developing rheumatoid arthritis. 20% of the patients with rheumatoid arthritis do not have a susceptibility epitope, so this result alone does not exclude the diagnosis. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-22 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-15 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | DRB1*07:01, DRB1*07:01 | DRB1*04:01/04:04/04:05. DRB1*01:01, DRB1*14:02 to a lesser extent | This patient does not have a HLA-DRB1 allele classically associated with Rheumatoid Arthritis (RA). | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-DRB1*07:01, 07:01 | HLA-DRB1*01:01, DRB1*01:02, DRB1*04:01, DRB1*04:04, DRB1*04:05, DRB1*04:08, DRB1*10:01, DRB1*14:02 | HLA-DRB1*01:01, DRB1*01:02, DRB1*04:01, DRB1*04:04, DRB1*04:05, DRB1*04:08, DRB1*10:01, DRB1*14:02 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | No
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1265 | DRB1*07 DRB1*07 | DRB1*01:01/02/04 ; DRB1*10 ; DRB1*14:02 DRB1*04:01/04/05/08 | ABSENCE of a susceptibility epitope carried by a molecule HLA-DRB1. The patient does not have an epitope associated with a risk of developing rheumatoid arthritis, but 20% of patients with rheumatoid arthritis do not have one. This result alone does not exclude the diagnosis of rheumatoid arthritis (ACR 2010 criteria). (Ref : Balandraud N, Picard C, Reviron D, Landais C, Toussirot E, Roudier J. HLA-DRB1 Genotypes and the Risk of Developing Anti- Citrullinated Protein Antibody (ACPA) Positive Rheumatoid Arthritis. PLoS One (2013) 8:64108) | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | No
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-10 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-1-2026 Diabetes
8-1-2026 | Diabetes
Distribution Date: 30/06/2026
Total distributed | 8
Total submitted | 52
Not tested | 1
Reference | DRB1*07:01, 07:01; DQA1*02:01, 02:01; DQB1*02:02, 03:03
Number acceptable | 7
Assessed | Yes
Comments
Lab Code | Diabetes Results for Assessment | Diabetes Alleles of Interest | Diabetes Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | No
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | HLA-DRB1*03 :01 ;DQA1*05 :01 ;DQB1*02 :01 ABSENT HLA-DRB1*04 :01 ;DQA1*03 :01 ;DQB1*03 :02/04 ABSENT HLA-DRB1*04 :02 ;DQA1*03 :01 ;DQB1*03 :02/04 ABSENT HLA-DRB1*04 :04 ;DQA1*03 :01 ;DQB1*03 :02/04 ABSENT HLA-DRB1*04 :05 ;DQA1*03 :01 ;DQB1*03 :02/04 ABSENT HLA-DRB1*15 :01 ; DQA1*01 :02 ; DQB1*06 :02 ABSENT HLA-DRB1*14 :01 ; DQA1*01 :01 ; DQB1*05 :03 ABSENT HLA-DRB1*07 :01 ; DQA1*02 :01 ; DQB1*03 :03 PRESENT HLA-DRB1*04 :03 ; DQA1*03 :01 ; DQB1*03 :02 ABSENT | HLA-DRB1*03 :01 ;DQA1*05 :01 ;DQB1*02 :01 HLA-DRB1*04 :01 ;DQA1*03 :01 ;DQB1*03 :02/04 HLA-DRB1*04 :02 ;DQA1*03 :01 ;DQB1*03 :02/04 HLA-DRB1*04 :04 ;DQA1*03 :01 ;DQB1*03 :02/04 HLA-DRB1*04 :05 ;DQA1*03 :01 ;DQB1*03 :02/04 HLA-DRB1*15 :01 ; DQA1*01 :02 ; DQB1*06 :02 HLA-DRB1*14 :01 ; DQA1*01 :01 ; DQB1*05 :03 HLA-DRB1*07 :01 ; DQA1*02 :01 ; DQB1*03 :03 HLA-DRB1*04 :03 ; DQA1*03 :01 ; DQB1*03 :02 | Présence d'allèles HLA-DRB1*07:01/DQA1*02:01/DQB1*03:03, pouvant constituer un haplotype HLA protecteur conférant à l'individu porteur un risque plus faible de développer la maladie. Ce résultat ne constitue pas un critère de diagnostic d'exclusion de la maladie. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | No tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | No
Lab-1038 | HLA genotype : HLA-DRB1*07:01, DRB1*- (homozygous) HLA-DQA1*02:01, DQA1*- (homozygous) HLA-DQB1*02:02, DQB1*03:03 Result : HLA-DRB1*03, DQA1*05:01, DQB1*02:01: absence HLA-DRB1*04:01/04:02/04:04/04:05, DQA1*03, DQB1*03:02 : absence HLA-DRB1*04:01/04:02/04:04/04:05, DQA1*03, DQB1*03:04 : absence | HLA-DRB1*03:01, DQA1*05:01, DQB1*02:01 HLA-DRB1*04:01/04:02/04:04/04:05, DQA1*03, DQB1*03:02 HLA-DRB1*04:01/04:02/04:04/04:05, DQA1*03, DQB1*03:04 | This individual has no HLA haplotype associated with type 1 diabetes. The DR3,DQ2 and DR4,DQ8 haplotypes are found in 95% of type 1 diabetes patients whereas the HLA-DR3 and DR4 antigens are found 40% of the Caucasian population. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | No
Lab-1079 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | DRB1*07:01,- | DRB1*03:01;DQA1*05:01;DQB1*02:01 DRB1*04:01;DQA1*03:01;DQB1*03:02/04 DRB1*04:02;DQA1*03:01;DQB1*03:02/04 DRB1*04:04;DQA1*03:01;DQB1*03:02/04 DRB1*04:05;DQA1*03:01;DQB1*03:02/04 | Absence of alleles that may constitute a susceptibility HLA haplotype to type I diabetes. This result does not constitute a criterion for excluding the disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-22 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-15 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | DQB1*02:02, DQB1*03:03 | DQB1*03:02, DQB1*02 | This patient does not have a HLA genotype associated with Insulin-Dependent Diabetes Mellitus (IDDM). | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-DRB1*07:01, 07:01; DQA1*02:01, 02:01; DQB1*02:02, 03:03 | HLA-DRB1*03:01/DQA1*05:01/DQB1*02:01; HLA-DRB1*0405/DQA1*03:01/03/DQB1*03:02; HLA-DRB1*04:01/DQA1*03:01/03/DQB1*03:02; HLA-DRB1*04:02/DQA1*03:01/DQB1*03:02 | HLA-DRB1*03:01/DQA1*05:01/DQB1*02:01; HLA-DRB1*0405/DQA1*03:01/03/DQB1*03:02; HLA-DRB1*04:01/DQA1*03:01/03/DQB1*03:02; HLA-DRB1*04:02/DQA1*03:01/DQB1*03:02 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | No
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | HLA-DRB1*07 | HLA-DRB1* |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1239 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1265 | DQB1*02:02 ; DQA1*02:01 ; DRB1*07 DQB1*03:03 ; DQA1*02:01 ; DRB1*07 | DQB1*02:01/DQA1*05:01/DRB1*03:01 ; DQB1*03:02/DQA1*03/DRB1*04:01/02/04/05 | Investigation of susceptibility to type I diabetes ABSENCE of an HLA haplotype of susceptibility to type I diabetes. This result does not constitute a criterion for exclusion of the disease. (Ref: Noble JA. Immunogenetics of type 1 diabetes: A comprehensive review. J Autoimmun. nov 2015;64:101-12. Zhao LP et al nine residues in HLA-DQ molecules determine with susceptibility and resistance to type 1 diabetes among young children in Sweden. Sci Rep . 2021 Apr 23;11(1):8821) | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | No
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-10 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-1-2026 Psoriasis
8-1-2026 | Psoriasis
Distribution Date: 30/06/2026
Total distributed | 6
Total submitted | 52
Not tested | 1
Reference | C*06:02, 16:01
Number acceptable | 5
Assessed | Yes
Comments
Lab Code | Psoriasis Results for Assessment | Psoriasis Alleles of Interest | Psoriasis Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | No
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | C*06:02/07/13/14/15++, *16:01/07/08/10/11++ | HLA-C*06 | This patient appears to have the HLA-C*06 allele group associated with psoriasis | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | No
Lab-1038 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | C*06 positive | C*06 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | No
Lab-1079 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-22 | No
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-15 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | B*44, B*57 C*06, C*16 | C*06 or B*27 | This patient has the HLAC*06 allele associated with skin involvement in psoriatic disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-C*06:02, 16:01 | HLA-C*06 | HLA-C*06 positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | No
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | Number of HLA-C*06:02 alleles: 1. SNP genotypes are rs10484554: CT, rs3130457: CT, rs6904246: GG, rs7745906: AG. |  |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1265 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | No
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-10 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## hiddenSheet (hidden)
vunhst_results_scheme8:JJWqPNKIjfrmJ23mEc3pI05q3aFHkp+qaKcIC8X7hXChZTueLk4axw8dvRfjk+uXcHAIfjHAcvaQnt0bnWeszA==:vunhst_results_scheme8id=%28Do%20Not%20Modify%29%20Results_Scheme8&checksumLogicalName=%28Do%20Not%20Modify%29%20Row%20Checksum&modifiedon=%28Do%20Not%20Modify%29%20Modified%20On&vunhst_sample=Sample&219de229-aa48-46be-87ca-da7db3f1735f.vunhst_laboratorycode=Laboratory%20Code%20%28Laboratory%29%20%28Laboratory%29&vunhst_laboratory=Laboratory&vunhst_coelresults=Coeliac%20Disease%20Results%20for%20Assessment&vunhst_coelalleles=Coeliac%20Disease%20Alleles%20of%20Interest&vunhst_coelintercomments=Coeliac%20Disease%20Interpretive%20Comments&vunhst_narcresults=Narcolepsy%20Results%20for%20Assessment&vunhst_narcalleles=Narcolepsy%20Alleles%20of%20Interest&vunhst_narcintercomments=Narcolepsy%20Interpretative%20Comments&vunhst_actiprurresults=Actinic%20Prurigo%20Results%20for%20Assessment&vunhst_actipruralleles=Actinic%20Prurigo%20Alleles%20of%20Interest&vunhst_actiprurintercomments=Actinic%20Prurigo%20Interpretative%20Comments&vunhst_birdretiresults=Birdshot%20Retinopathy%20Results%20for%20Assessment&vunhst_birdretialleles=Birdshot%20Retinopathy%20Alleles%20of%20Interest&vunhst_birdretiintercomments=Birdshot%20Retinopathy%20Interpretative%20Comments&vunhst_behcresults=Behcet%27s%20Disease%20Results%20for%20Assessments&vunhst_behcalleles=Behcet%27s%20Disease%20Alleles%20of%20Interest&vunhst_behceintercomments=Behcet%27s%20Disease%20Interpretative%20Comments&vunhst_rheuarthresults=Rheumatoid%20Arthritis%20Results%20for%20Assessment&vunhst_rheuarthalleles=Rheumatoid%20Arthritis%20Alleles%20of%20Interest&vunhst_rheuarthintercomments=Rheumatoid%20Arthritis%20Interpretative%20Comments&vunhst_diabresults=Diabetes%20Results%20for%20Assessment&vunhst_diaballeles=Diabetes%20Alleles%20of%20Interest&vunhst_diabintercomments=Diabetes%20Interpretative%20Comments&vunhst_psorresults=Psoriasis%20Results%20for%20Assessment&vunhst_psoralleles=Psoriasis%20Alleles%20of%20Interest&vunhst_psorintercomments=Psoriasis%20Interpretative%20Comments&vunhst_issubmitted=isSubmitted&vunhst_tested_new=Tested&vunhst_reasonfornottesting=Reason%20for%20not%20testing&vunhst_confirmreason=Confirm%20Reason&vunhst_datereceived=Date%20Received%20&vunhst_datetested=Date%20Tested
Yes | No