## 8-3-2026 Coeliac
8-3-2026 | Coeliac Disease
Distribution Date: 30/06/2026
Total distributed | 54
Total submitted | 53
Not tested | 1
Reference | DRB1*03:01, 04:07; DQA1*03:03, 05:01; DQB1*02:01, 03:01
Number acceptable | 50
Assessed | Yes
Comments
Lab Code | Coeliac Disease Results for Assessment | Coeliac Disease Alleles of Interest | Coeliac Disease Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-DQA1*03:03, 05:01; HLA-DQB1*02:01, 03:01 | Specific genotype combinations of HLA-DQA1 and DQB1 alleles (HLA-DQA1*05, HLA-DQB1*02 / HLA-DQA1*03, HLA-DQB1*03:02 / HLA-DQA1*02:01, HLA-DQB1*02:02) | Positive for HLA-DQ2.5 The patient carries the HLA-DQ2.5 heterodimer (HLA-DQA1*05, HLA-DQB1*02) associated with Coeliac disease. The following considerations should be made when interpreting the HLA status of the patient: • Major histocompatibility antigens are expressed for life. Approximately 95% of coeliac disease (CD) patients express the HLA-DQ heterodimers DQ2.5 (encoded by alleles DQB1*02 and DQA1*05) and DQ8 (encoded by alleles DQA1*03 and DQB1*03(DQ8)). • Expression of HLA-DQ2 and/or -DQ8 molecules is necessary, but not sufficient, to cause the disease. This result must be used in combination with further clinical information. • HLA-DQ2 and -DQ8 is present in approximately 40% of the UK population so a positive test in the absence of CD specific antibody testing has no predictive value, but a negative test can exclude a diagnosis of CD (NICE Clinical Guideline 86 2009). • Whilst HLA typing does not have an absolute diagnostic value it enables assessment of relative risk in the presence of a positive CD antibody test. References: Bodd et al. (2012) T-cell response to gluten in patients with HLA-DQ2.2 reveals requirement of peptide-MHC stability in celiac disease. Gastroenterology 142:552-561. D. Pritchard et al. (2024) UK NEQAS and BSHI guideline: Laboratory testing and clinical interpretation of HLA genotyping results supporting the diagnosis of coeliac disease. International Journal of Immunogenetics 51:3-20. Medrano et al. (2012) HLA and celiac disease susceptibility: new genetic factors bring open questions about the HLA influence and gene-dosage effects. PLoS ONE 7(10):e48403. Megiorni et al. (2009) HLA-DQ status and risk gradient for celiac disease. Human Immunology 70:55-59. | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | HLA-DQ2.5 positive. | DQA1*02, DQA1*03, DQA1*05 DQB1*02, DQB1*03:02 | The patient is HLA-DQ2.5 positive. Celiac disease is associated with this type of HLA that that occurs in 95% of the cases. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1008 | DQA1*03, *05; DQB1*02:01/05/07/08/09++, *03:01/19/21/22/24++ | HLA-DQA1*05/DQB1*02 HLA-DQA1*03/DQB1*03:02 HLA-DQA1*02/DQB1*02 | This person has the HLA-DQB1*02/DQA1*05 allele combination encoding the coeliac-associated HLA-DQ2 heterodimer | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | DQA1*02 not detected, DQA1*03 detected, DQA1*05 detected, DQB1*02 detected, DQB1*03:02 not detected | DQA1*02, DQA1*03, DQA1*05, DQB1*02, DQB1*03:02 | Subject has the HLA-DQ2.5 allotype which is expressed by >85% of patients with coeliac disease. However, the presence of associated HLA alleles should not be used to diagnose coeliac disease due to the very low positive predictive value of the association. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1020 | DQA1*05, DQB1*02 | DQ2.5 (DQA1*05, DQB1*02) DQ2.2 (DQA1*02, DQB1*02) DQ8 (DQA1*03, DQB1*03:02) |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1024 | HLA-DQA1*02 ABSENT HLA-DQA1*05 PRESENT HLA-DQB1*02 PRESENT HLA-DQB1*03:02 (DQ8) ABSENT | HLA-DQA1*02,HLA-DQA1*05,HLA-DQB1*02 and HLA-DQB1*03:02 (DQ8) | Présence du HLA-DQ2.5. Risque élevé de prédisposition à la maladie coeliaque. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1025 | GENETIC PREDISPOSITION DETECTED | HLA-DQA1*05 HLA-DQB1*02 (the kit also detected the HLA-DQB1-no*02 allele, indicating a heterozygous genotype at this locus) HLA-DQB1*03:01 HLA-DQA1*03 HLA-DRB1*03 HLA-DRB1*04 | The detected alleles are compatible with expression of the HLA-DQ2.5 heterodimer, which is associated with genetic susceptibility to coeliac disease. HLA typing identifies genetic predisposition only and does not establish the diagnosis of coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | Yes
Lab-1026 | DQA1*05=positive; DQB1*02= positive; DQB1*03:02/05 (DQ8 group)= negative | DQA1*05; DQB1*02; DQB1*03:02/05 (DQ8 group) |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1027 | The sample contains HLA-DQ2.5 (HLA-DQA1*05 and HLA-DQB1*02), genetic marker associated with celiac disease. |  |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | Yes
Lab-1031 | DQA1*05= positive; DQB1*02=positive; DQB1*03:02=negative | DQA1*05, DQB1*02, DQB1*03:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | Yes
Lab-1033 | HLA-DQB1*02: positive, HLA-DQA1*05:01: positive, HLA-DQB1*03:02: negative | HLA-DQA1*05:xx, HLA-DQB1*02:01,02, HLA-DQB1*03:02 | The allele DQ2.5, which is strongly associated with coeliac disease, could be detected. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | Yes
Lab-1038 | HLA genotype : HLA-DQA1*05:01, DQA1*03:03 HLA-DQB1*02:01, DQB1*03:01 Result : HLA-DQ2.5 (DQA1*05:01, DQB1*02:01, cis): presence HLA-DQ2.5 (DQA1*05:05, DQB1*02:02, trans) : absence HLA-DQ2.2 (DQA1*02:01, DQB1*02:02, cis) : absence HLA-DQ8 (DQA1*03, DQB1*03:02) : absence | HLA-DQ2.5 (DQA1*05:01, DQB1*02:01, cis) HLA-DQ2.5 (DQA1*05:05, DQB1*02:02, trans) HLA-DQ2.2 (DQA1*02:01, DQB1*02:02, cis) HLA-DQ8 (DQA1*03, DQB1*03:02) | This individual carries HLA-DQA1*05:01, DQB1*02:01 (DQ2.5) that has a strong association with coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest coeliac disease. >95% of CD patients carry HLA-DQ2 encoded by DQ2.5 (cis or trans) and 5% of CD patients carry HLA-DQ8 encoded by DQA1*03, DQB1*03:02. HLA genotyping has a good negative predictive value in the diagnosis of CD. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | Positive association for Coeliac disease: DQA1*03:03-DQB1*03:01 and DQA1*05:01-DQB1*02:01 | DQA1 & DQB1 (DQ2/DQ8) | Presence of DQ2.5cis and DQ8 | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | Yes
Lab-1057 | DQA1*03,DQA1*05,DQB1*02,DQB1*03:01 | DQA1*05,DQB1*02,DQB1*03:02 | DQA1*05=present DQB1*02=present DQB1*03:02=absent | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | HLA-DQA1*03:03/HLA-DQA1*05:01 HLA-DQB1*02:XX/HLA-DQB1*03:XX (DQ7) |  |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | Yes
Lab-1065 | DQB1*02 heterozygous DQA1*05 positive DQB1*03:02 negative DQA1*03 positive | DQB1*02 (heterozygous/homozygous), DQA1*05, DQB1*03:02 and DQA1*03 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | DQ2.5: Positive DQ2.2: Negative DQ8: Negative | Combination of: DQ2.5: DQA1*05 and DQB1*02 DQ2.2: DQA1*02 and DQB1*02 DQ8: DQA1*03 and DQB1*0302 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1077 | DQA1*03:03:01, DQA1*05:01:01, DQB1*02:01, DQB1*03:01 | DQA1 and DQB1 | DQB1*02, DQA1*05 (DQ2.5) Positive DQB1*02, DQA1*02 (DQ2.2) Negative DQB1*03:02 (DQ8) Negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | Yes
Lab-1079 | DQB1*02:01:01 DQB1*03:01:01 DQA1*03:03:01 DQA1*05:01:01 | HLA-DQB1*02, HLA-DQB1*03:02 and HLA-DQA1*05 | This patient is POSITIVE for HLA-DQ2 (DQA1*05, DQB1*02) and NEGATIVE for HLA-DQ8. Patients with this genotype have a HIGH RISK of predisposition to Coeliac disease though other factors are likely involved | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | DQB1*02,*03 (no *03:02) DQA1*03,*05 | DQA1, DQB1 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | Yes
Lab-1085 | DQ2.2 negative, DQ2.5 positive, DQ8 negative | DQA1*02 DQA1*03 DQA1*05 DQA1*01/04/06 DQA1*02/03:01 DQA1*03:02/03 DQB1*02 DQB1*03:02 DQB1*03/06 DQB1*04 DGB1*05 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1086 | HLA DQ2.5: positive HLA DQ2.2: negative HLA DQ8: negative | DQA1*05, DQB1*02, DQA1*02, DQA1*03, DQB1*03, DQB1*0302 *0305, DQB1*0302 *0303 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | Yes
Lab-1090 | DQA1*03,*05 DQB1*02,*03(serological equivalent DQ7) | DQB1*02+DQA1*05 DQB1*02 DQB1*03:02 | Presence of the HLA-DQB1*02/HLA-DQA1*05 alleles, associated with a high risk of developing celiac disease, higher than that of the general population. This result is in favor of celiac disease only if it is associated with the clinical, biological and/or endoscopic signs corresponding to the diagnostic criteria. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-14 | Yes
Lab-1104 | HLA coeliac disease associated Heterodimer result: HLA-DQB1*02, DQA1*05 (DQ2.5) DETECTED HLA-DQB1*02, DQA1*02 (DQ2.2) NOT detected HLA-DQB1*03:02 (DQ8) NOT detected | HLA Coeliac Disease associated Heterodimers: HLA-DQB1*02, DQA1*05 (DQ2.5) HLA-DQB1*02, DQA1*02 (DQ2.2) HLA-DQB1*03:02 (DQ8) | This individual has a genotype which is associated with coeliac disease (CD). More than 90% of individuals with coeliac disease (CD), across all populations, express either HLA-DQ2 and / or DQ8 molecules with less than 2% lacking these human leukocyte antigen (HLA) types completely. The HLA-DQA1 and DQB1 alleles that encode these CD associated heterodimers (also referred to as DQ2.5, DQ2.2 and DQ8) are DQA1*05 DQB1*02, DQA1*02 DQB1*02 and DQB1*03:02. The majority of individuals with CD express the DQ2.5 heterodimer while remaining 5-10% express DQ2.2 or DQ8. As HLA-DQ2 and DQ8 are common in the general population (40% in Europeans) and only 1% develop CD, their presence has a low positive predictive value and does not confer a diagnosis of CD. A diagnosis should be based on serological testing from the gluten sensitivity profile, clinical findings and histological evaluation of duodenal biopsy. For further information see Pritchard, D. et al., 2024. UK NEQAS and BSHI guideline: Laboratory testing and clinical interpretation of HLA genotyping results supporting the diagnosis of coeliac disease. International Journal of Immunogenetics, 51(Suppl. 1), 3–20. This assay has been performed by PCR-SSO (sequence specific oligonucleotides) based on Luminex xMAP technology using the LIFECODES HLA-DQA1/B1 SSO Typing kit (CE IVD) (Immucor), to determine the presence or absence of specific HLA alleles that form the above listed CD associated heterodimers. Heterodimer ambiguities have been resolved using additional PCR-SSP (sequence specific primers with the HLA Coeliac Disease Domino System (CE IVD) (Protrans), kit lot number 13-0 version 01 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1111 | HLA-DQ2.5 - positive (+) HLA-DQ2.2 - negative (-) HLA-DQ8 - negative (-) | HLA-DQ2.5 (DQA1*05/DQB1*02), HLA-DQ2.2 (DQA1*02/DQB1*02), HLA-DQ8 (DQA1*03/DQB1*03:02) | Haplotype: DQ2.5 Disease risk: high Only a negative result for the HLA-DQ2/DQ8 alleles rules out a genetic basis for celiac disease with 99% probability. A positive result does not confirm the disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1118 | DQB1*02:01, DQB1*03:01 DQA1*03:03, DQA1*05:01 | DQB1*02:01, DQB1*02:02, DQB1*03:02, DQA1*05 | The patient possesses the following HLA alleles that are associated with Coeliac Disease: DQA1*05 and DQB1*02. Patients with this genotype have a high risk of predisposition to Coeliac Disease, however; presence of these alleles alone does not confirm diagnoses. Other clinical indications are required for diagnosis. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1120 | • DQA1*05 Detected • DQB1*02 Detected • DQB1*0302 Not detected | DQA1*05, DQB1*02, DQB1*0302 | This individual has one of the HLA-DQ variants associated with coeliac disease (DQ2.5). More than 97% of coeliac disease patients carry either HLA-DQ2.2, HLA-DQ2.5 or DQ8. However, these variants are also present in approximately 40% of the general population and therefore whilst possession of the variant can support a diagnosis of coeliac disease it is not per se diagnostic of the condition. This assay tests for the presence of HLA DQ2.2 (DQB1*02), HLA-DQ2.5 (DQA1*05/DQB1*02) and HLA-DQ8 (DQB1*0302). | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | Yes
Lab-1132 | DQA105 : positive; DQB102: positive; DQB1*03:02: negative | DQA105 - DQB102 - DQB1*03:02 This test cannot detect variants other than those targeted by the analysis. Test sensitivity and specificity: 99%. | DQ2.5 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | Yes
Lab-1134 | HLA-DQB1*02 positive, HLA-DQA1*05 positive, HLA-DQB1*03:02 negative | HLA-DQB1*02, HLA-DQB1*03:02, HLA-DQA1*05 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1139 | DQA1*03:03, DQA1*05:01 DQB1*02:01+DQB1*03:01 | DQA1*05 + DQB1*02 (cis or trans) DQA1*XX + DQB1*03:02 DQA1*02 + DQB1*02 | The presence of DQA1*05, DQB1*02 (DQ2.5) has a strong association with coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest coeliac disease. The presence of an associated HLA genotype does not confer a diagnosis of coeliac disease and has a low positive predictive value for coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | DQB1*03:02: negative DQA1*05: positive DQB1*02: positive DQA1*02: negative. Reported serotypes: DQ2.5 | HLA-DQB1*03:02, HLA-DQA1*05, HLA-DQB1*02, HLA-DQA1*02 | "HLA-DQ2.5 is detected in the form of HLA-DQB1*02 and HLA-DQA1*05. Most coeliac patients have these alleles. The alleles are common in the general population. Positive test alone is not sufficient to diagnose coeliac disease but makes the diagnosis more probable." | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | Yes
Lab-1146 | DQA1*05, DQA1 DQB1*02, DQB1 | DQA1*05 DQB1*02, DQB1*03:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1150 | DQ2.5 | DQA1*05-DQB1*02-DRB1*03 DQA1*05-DQB1*03:01-DRB1*11/DRB1*12 DQA1*02-DQB1*02-DRB1*07 DQA1*03-DQB1*03:02-DRB1*04 | detected HLA-genotype is associated with a risk of coeliac disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | Yes
Lab-1151 | DQA1*03:03, 05:01 DQB1*02:01, 03:01 | DQA1*05 , DQA1*02, DQB1*02, DQB1*03:02 | This presence ofDQA1*05, DQB1*02 (HLA-DQ2.5) has a strong association with coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | Yes
Lab-1154 | DQ2 present | DQ2, DQ8 and DQ7 and DQA05 | DQ2 is associated with Coeliac disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1176 | HLA-DQA1*03:03, 05:01; DQB1*02:01, 03:01 | HLA-DQA1*05:01/DQB1*02:01 (DQ2.5); HLA-DQA1*03/DQB1*03:02 (DQ8); HLA-DQA1*02:01/DQB1*02:02 (DQ2.2) | Positive for DQ2.5. The presence of DQA1*05, DQB1*02 (HLA-DQ2.5) has a strong association with coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-08-01 | 2026-08-03 | Yes
Lab-1186 | DQA1*05 - POSITIVE DQB1*02 - POSITIVE, Heterozygosity DQA1*03 - POSITIVE DQB1*03:02 - NEGATIVE DQB1*03:03 - NEGATIVE | HLA-DQA1*05, HLA-DQA1*03, HLA-DQB1*02, HLA-DQB1*03:02, HLA-DQB1*03:03 | DQ2 haplotype. Risk of celiac disease | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1199 | DQB1*02-DQA1*05 DQB1*03:01-DQA1*03 | HLA-DQ2.5/DQ8 | This presence of HLA-DQA1*05/DQB1*02 (DQ2.5) has a strong association with coeliac disease in patients where laboratory tests or symptoms or endoscopic features suggest coeliac disease. The presence of an associated HLA genotype does not confer a diagnosis of coeliac disease and has a low positive predictive value for coeliac disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | HLA DQ2.2 = Absent ; HLA DQ2.5 = Present ; HLA DQ8 = Absent | DQA1*02; DQA1*02/*0301; DQA1*03; DQA1*0302/03; DQA1*05; DQB1*02; DQB1*02/*0302. |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | Genotypen er associeret med cøliaki. Analysen har påvist HLA-DQB1*02/HLA-DQA1*05. Den genetiske risiko for cøliaki er til stede. | DQA1*05 positive DQB1*02 positive DQA1*02 negative DQB1*03:02 negative DQA1*03 positive |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | Yes
Lab-1208 | DQB1*02 - Positive DQA1*05 - Positive DQB1*03:02 - Negative | DQB1*02 DQA1*05 DQB1*03:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | Yes
Lab-1209 | HLA-DQ2.2 - NEGATIVE HLA-DQ2.5 - POSITIVE HLA-DQ8 - NEGATIVE b-subunit HLA-DQ2.2/DQ2.5 - POSITIVE |  |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | Yes
Lab-1215 | DQB1*03:02 - Negative DQB1*02 - Positive DQA1*05 - Positive | DQB1*03:02 DQB1*02 DQA1*05 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1239 | This patient is POSITIVE for the HLA-DQ2.5 genotype but negative for HLA-DQ8 and -DQ2.2. | DQA1*02, *03, *05 and DQB1*02, *03:02 | CD is associated with HLA heterodimers HLA-DQA1*05-DQB1*02 (DQ2.5), HLA-DQA1*03-DQB1*03:02 (DQ8) or HLA-DQA1*02-DQB1*02 (DQ2.2). The absence of any of these associated genotypes excludes the diagnosis of coeliac disease with a high probability (<1% likelihood). Further information is available in the ESsCD guidelines (2019): https://doi.org/10.1177/2050640619844125. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1265 | DQB1*02:01 ; DQA1*05:01 DQB1*03:01 ; DQA1*03:02 | DQB1*02:01/DQA1*05:01 ; DQB1*03:02/DQA1*03 DQB1*02:02/DQA1*02:01 | Investigation of susceptibility to celiac disease Presence of the HLA-DQB1*02:01/HLA-DQA1*05:01 (DQ2.5) haplotype This haplotype is associated with a significantly increased risk of developing celiac disease compared to the general population. However, this result supports a diagnosis of celiac disease only when correlated with a clinical presentation and other laboratory data, according to established diagnostic criteria. Absence of the HLA-DQB1*03:02/HLA-DQA1*03 (DQ8) haplotype (Réf. : Ludvig M.Sollid et al., 2012 ; Megiorni F et al. 2012 ; Hadithi M. et al. 2007 ; Margaritte-Jeannin P et al., 2004) | Unacceptable | Reported DQA1*03:02 instead of DQA1*03:03 | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1268 |  |  |  | Unacceptable | No results returned | No | [5 empty] | Yes
Lab-1269 | Presence of DQB1*02, DQA1*05, DQA1*03, absence of DQB1*03:02 | DQB1*02, DQA1*05, DQB1*03:02, DQA1*03 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | Yes
Lab-1272 | HLA-DQA1*05 -> positive HLA-DQB1*02 -> positive HLA-DQB1*03:02 -> negative | HLA-DQA1*05, HLA-DQB1*02, HLA-DQB1*03:02 | DQ2.5 +; DQ8 - Hohes Risiko Glutenunverträglichkeit | Acceptable |  | Yes | Partially Tested | Other | Only coeliac disease is tested in the laboratory. No other analyses are performed. | 2026-07-06 | 2026-08-04 | Yes
Lab-1284 | DQB1*02 – DQA1*05 | HLA-DQB1*03:02; HLA-DQB1*02 – HLA-DQA1*05; HLA-DQB1*02 – HLA-DQA1*02 | Positive for HLA-DQ2.5, which is associated with coeliac disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | Yes
Lab-1291 | DRB1*03 - DQA1*05 - DQB1*02 DRB1*04 - DQA1*03 - DQB1*03:01 |  | DR3 DQ2 DR4 DQ7 | Acceptable |  | Yes | Partially Tested | Other | Our laboratory tests only HLA for Coeliac Desease | 2026-07-06 | 2026-07-08 | Yes
Lab-1295 | Identified heterodimers: DQ2.5 cis + DQ8 | // | // | Unacceptable | Reported DQ8 false positive | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | Yes
Lab-1296 | DQA1*05 DQB1*02 (DRB1*03, DQB1*0301, DQA1*03, DRB1*04) | DQA1*05 DQB1*02 DRB1*03 DQB1*0301 DQA1*03 DRB1*04 | HLA-DQ2.5 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1304 | HLA-DQA1*05 Positive HLA-DQB1*02 Positive/Heterozygous HLA-DQB1*03:02 Negative | HLA-DQA1*05 HLA-DQB1*02 HLA-DQB1*03:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | Yes
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-3-2026 Narcolepsy
8-3-2026 | Narcolepsy
Distribution Date: 30/06/2026
Total distributed | 25
Total submitted | 52
Not tested | 1
Reference | DQB1*02:01, 03:01
Number acceptable | 24
Assessed | Yes
Comments
Lab Code | Narcolepsy Results for Assessment | Narcolepsy Alleles of Interest | Narcolepsy Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-DQB1*02:01, 03:01 | HLA-DQB1*06:02 | This patient is Negative for HLA-DQB1*06:02, the allele associated with Narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | DQB1*06:02 negative | DQB1*06:02 | The patient is lacking the HLA-type (DQB1*06:02) wich is associatetd with narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1008 | DQB1*02:01/05/07/08/09++, *03:01/19/21/22/24++ | HLA-DQB1*06:02 | This patient does not have the HLA-DQB1*06:02 allele associated with narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | We didn't find narcolepsy associated alleles in sample 03. | DQA1*01:02, DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1024 | DQB1*06:02 ABSENT | DQB1*06:02 | Absence de l'allèle recherche HLA-DQB1*06:02. Le diagnostic de narcolepsie de type 1 est peu probable mais n’est pas formellement exclu. L’interprétation de ce résultat doit être réalisée en fonction du contexte clinique. Références : - Rennac V et al. Guidelines from the French-speaking Society for Histocompatibility and Immunogenetics (SFHI) for harmonisation of HLA genotyping in autoimmune diseases, drug hypersensitivity and pharmacogenetics. HLA 2025;106:e70442. - Capittini C et al. Correlation between HLA-DQB1*06 :02 and narcolepsy with and without cataplexy. Sleep Medicine 2018;52:150-157. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | DQB1*06:02=negative | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | HLA-DQB1*06:02: negative | HLA-DQB1*06:02 | In absence of the Allele HLA-DQB1*06:02, narcolepsy type I with cataplexy can be excluded with a sensitivity of approximately 90%. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | Yes
Lab-1038 | HLA-DQB1*02:01, DQB1*03:01 | HLA-DQB1*06:02 | This individual has no DQB1* allele associated with narcolepsy. The HLA-DQB1*06:02 allele is found in 15-25% of the overall population and in 90-100% of narcolepsy patients. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | No known association for narcolepsy | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | Yes
Lab-1057 | DQB1*02,DQB1*03 | DQB1*06:02 | DQB1*06:02=absent | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | DQB1*06:02 negative | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | DQB1*02:01, DQB1*03:01 | DQB1*06:02 | DQB1*06:02 Negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | Yes
Lab-1079 | DQB1*02:01:01 DQB1*03:01:01 | HLA-DQB1*06:02 | HLA-DQB1*06:02 Negative. Narcolepsy is associated with the expression of the human leukocyte antigen (HLA) class II molecule DQB1*06:02. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | DQB1*02,*03(serological equivalent DQ7) | DQB1*06:02 | Absence of the susceptibility allele for narcolepsy-cataplexy DQB1*06:02. This allele is present in 85 to 95% of the patients with narcolepsy-cataplexy, in 40 to 60% of patients with narcolepsy without cataplexy and in 18% of patients with idiopathic hypersomnia but it is also present in 12 to 38% of the population. This result makes the diagnosis of narcolepsy-cataplexy unlikely but does not exclude the diagnosis. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-14 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | DQB1*02:01, DQB1*03:01 | DQB1*06:02 | HLA-DQB1*06:02 is associated with narcolepsy-cataplexy. This patient is NEGATIVE for HLA-DQB1*06:02. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | HLA-DQB1*06:02 negative | HLA-DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1139 | DQB1*02:01, DQB1*03:01 | DQB1*06:02 | This patient does not have the HLADQB1*06 allele classically associated with Narcolepsy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-DQB1*02:01, 03:01 | HLA-DQB1*06:02 | HLA-DQB1*06:02 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-08-01 | 2026-08-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | DQB1*02-DQA1*05 DQB1*03:01-DQA1*03 | HLA-DQB*06:02 | Patient has tested negative for HLA DQB1*06:02 allele which is associated with Narcolepsy. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | DQB1*06:02 - Negative | DQB1*06:02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | Yes
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | HLA-DQB1*06:02 NEGATIVE | DQB1*06:02 | The HLA-DQB1*06:02 allele associated with, but not diagnostic of, Narcolepsy is absent in this patient. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1265 | DQB1*02:01 DQB1*03:01 | DQB1*06:02 | Investigation of Susceptibility to Narcolepsy ABSENCE of the HLA-DQB1*06:02 susceptibility allele described in narcolepsy-cataplexy. This result makes the diagnosis of narcolepsy-cataplexy unlikely, but does not exclude the diagnosis. It should be compared with the clinical and biological signs and the functional tests corresponding to the diagnostic criteria. (Réf : Correlation between HLA-DQB1*06:02 and narcolepsy with and without cataplexy: approving a safe and sensitive genetic test in four major ethnic groups. A Systematic meta analysis, Sleep Medecine, 2018) | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | Absence of DQB1*06 :02 | DQB1*06 :02 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | Yes
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | DQB1*06:02 negative. | HLA-DQB1*06:02 | Genotype is not associated with narcolepsy. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | Yes
Lab-1291 | [5 empty] | Yes | Partially Tested | Other | Our laboratory tests only HLA for Coeliac Desease | 2026-07-06 | 2026-07-08 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-3-2026 Actinic Prurigo
8-3-2026 | Actinic Prurigo
Distribution Date: 30/06/2026
Total distributed | 6
Total submitted | 53
Not tested | 0
Reference | DRB1*03:01, 04:07
Number acceptable | 6
Assessed | Yes
Comments
Lab Code | Actinic Prurigo Results for Assessment | Actinic Prurigo Alleles of Interest | Actinic Prurigo Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-DRB1*03:01, 04:07 | HLA-DRB1*04:07 | This patient is Positive for HLA-DRB1*04:07, the allele associated with Actinic Prurigo | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | DRB1*03:01/06/13/28/37++, *04:07/92/172/197/199++ | HLA-DRB1*04:07 | This patient is positive for the HLA-DRB1*04 allelic group and appears to have the HLA-DRB1*04:07 allele associated with Actinic Prurigo | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | No
Lab-1038 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1067 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | No
Lab-1079 | DRB1*03:01/03:06/03:13/03:28/03:37/03:45/03:46/03:47/03:48/03:50/03:51/03:55/03:56/03:58/03:59/03:61/03:62/03:67N/03:68N/03:72/03:73/03:80/03:82/03:83/03:98/03:100/03:104/03:106/03:107/03:109/03:110/03:114/03:116/03:117/03:121/03:123/03:124/03:127/03:129/03:132/03:134/03:136/03:137/03:138/03:139/03:142/03:143/03:144/03:146/03:147/03:149/03:151/03:152/03:153/03:156N/03:161/03:163/03:164/03:165/03:166/03:168/03:169/03:170/03:175/03:180/03:182/03:183/03:184/03:186/03:187/03:189N/03:192/03:193/03:194/03:195/03:196/03:197/03:198/03:199/03:201/03:202/03:203/03:204N/03:208/03:209/03:210/03:214/03:216/03:217/03:220/03:221/03:223N/03:224 DRB1*04:07/92/172/197/199/209/222/246/249/256/262/263/269/278/298/315/337/344/351/356/366/393/394 | HLA-DRB1*04:07 | HLA-DRB1*04:07 Positive. Actinic Prurigo is associated with the expression of the human leukocyte antigen (HLA) class II molecule DRB1*04:07 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-14 | No
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | DRB1*03:01, DRB1*04:07 | DRB1*04:07 | This patient has the DRB1*04:07 allele associated with Actinic Prurigo | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-DRB1*03:01, 04:07 | HLA-DRB1*04:07 | HLA-DRB1*04:07 positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-08-01 | 2026-08-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | DRB1*03 DRB1*04:07 | HLA-DRB1*04:07 | Positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1265 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | No
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Partially Tested | Other | Our laboratory tests only HLA for Coeliac Desease | 2026-07-06 | 2026-07-08 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | [5 empty] | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | No

## 8-3-2026 Birdshot
8-3-2026 | Birdshot
Distribution Date: 30/06/2026
Total distributed | 16
Total submitted | 53
Not tested | 0
Reference | A*01:01, 32:01
Number acceptable | 16
Assessed | Yes
Comments
Lab Code | Birdshot Results for Assessment | Birdshot Alleles of Interest | Birdshot Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-A*01:01, 32:01 | HLA-A29 | This patient is Negative for HLA-A29, the specificity associated with Birdshot Retinopathy | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | A*01:01/04N/09/11N/17++, *32:01/05/11Q/12/14++ | HLA-A*29 | This patient does not have the HLA-A29 antigen associated with birdshot retinopathy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | HLA-A*29: negative | HLA-A*29 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | Yes
Lab-1038 | HLA Genotype: HLA-A*01:01, A*32:01 Result: HLA-A*29 : absence | HLA-A*29 | The absence of the HLA-A*29 allele argues against the diagnosis of Birdshot chorioretinitis. The prevalence of HLA-A29 in patients with Birdshot chorioretinitis is 90–100% according to published studies. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | A*01,A*32 | A*29 | A*29=absent | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | A*29 negative | A*29 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | A*01:01:01, A*32:01:01 | A*29 | A*29 Negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | Yes
Lab-1079 | A*01:01:01 A*32:01:01 | HLA-A*29 | HLA-A*29 Negative. Birdshot retinochoroidopathy is associated with the expression of the human leukocyte antigen (HLA) class I molecule A*29 | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | negative | HLA-A29 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | A*01,*32 | A*29 | Absence of susceptibility allele A*29. In case of clinical signs of posterior uveitis, the absence of the HLA-A*29 allele is very little in favor of the diagnosis of birdshot disease. The prevalence of HLA-A*29 in birdshot patients is 90 to 100% according to published studies. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-14 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | A*01:01, A*32:01 | A*29 | The patient is negative for HLA-A29 which is associated with Birdshot chorioretinopathy. The absence of HLA-A29 alone cannot be used to eliminate disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | A*01, A*32 | A*29 | This patient does not have a HLA genotype associated with Birdshot Retinopathy | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-A*01:01, 32:01 | HLA-A*29 | HLA-A*29 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-08-01 | 2026-08-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | A*01 A*32 | HLA-A*29 | Negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | HLA-A29 NEGATIVE | HLA-A*29 | The HLA-A29 genotype associated with, but not diagnostic of, Birdshot Chorioretinopathy is absent in this patient. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1265 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | Absence of A*29 | A*29 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | Yes
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Partially Tested | Other | Our laboratory tests only HLA for Coeliac Desease | 2026-07-06 | 2026-07-08 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | [5 empty] | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | No

## 8-3-2026 Behcets
8-3-2026 | Behcets
Distribution Date: 30/06/2026
Total distributed | 20
Total submitted | 52
Not tested | 1
Reference | B*08:01, 51:01
Number acceptable | 19
Assessed | Yes
Comments
Lab Code | Behcets Results for Assessment | Behcets Alleles of Interest | Behcets Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | HLA-B*08:01, 51:01 | HLA-B51 | This patient is Positive for HLA-B51, the specificity associated with Behcet's Disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | Yes
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | B*08:01/05/08N/10/18++, *51:01/03/09/11N/14++ | HLA-B*51 | This patient appears to have the HLA-B51 antigen associated with Behcet’s disease | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | HLA-B*51 PRESENT | HLA-B*51 | « La présence de cet allèle confère au porteur un risque multiplié par 6 de développer une maladie de Behçet par rapport à la population générale. Il a été établi qu’environ 50 à 70% des patients atteintes possèdent cet allèle. Ce résultat ne permet pas à lui seul d’affirmer le diagnostic de maladie de Behçet. Il doit être associé aux autres signes clinico-biologiques correspondant aux critères diagnostiques. Aucune mesure préventive particulière ni recherche familiale ou anténatale ne sont requises chez un porteur de ce gène de prédisposition. » | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | HLA-B*51: positive | HLA-B*51 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | Yes
Lab-1038 | HLA genotype : HLA-B*08:01, B*51:01 Result : HLA-B*51: presence | HLA-B*51 | The patient carries the HLA-B*51 susceptibility allele. The presence of the HLA-B*51 allele confers a sixfold increased risk of developing Behçet's disease compared to the general population. It has been established that approximately 50% to 70% of patients carry this allele. However, this finding alone is insufficient to confirm a diagnosis of Behçet's disease; it must be considered alongside clinical signs that meet the diagnostic criteria. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | B*08,B*51 | B*51 | B*51=present | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | B*51 positive | B*51 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | B*08:01:01, B*51:01:01 | B*51 | B*51 Positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | Yes
Lab-1079 | B*08:01:01 B*51:01:01 | HLA-B*51, HLA-B*27 | HLA-B*51 Positive, HLA-B*27 Negative. Behcet's disease is associated with the expression of the human leukocyte antigen (HLA) class I molecule B51. HLA-B*27 result can be used for differential diagnoses due to the overlapping clinical features of Behcet's and HLA-B27 related diseases. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | Yes
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | positive | HLA-B51 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | Yes
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | B*08,*51 | B*51 | Presence of the susceptibility allele HLA-B*51. The presence of this allele confers on the wearer a risk multiplied by 6 to develop a Behçet disease compared to the general population. 30-60% of affected patients have this allele. This result alone does not support the diagnosis of Behçet's disease. It must be associated with other clinico-biological signs corresponding to the diagnostic criteria. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-14 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | B*08:01, B*51:01 | B*51. | The patient is HLA - B51 positive which is associated with Behcet’s disease. The presence of HLA - B51 alone is insufficient for diagnosis. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | B*08:01, B*51:01 | B*51 | This patient has a HLA genotype associated with Behcet's Disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-B*08:01, 51:01 | HLA-B51 | HLA-B51 positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-08-01 | 2026-08-03 | Yes
Lab-1186 | B*51 - POSITIVE | HLA-B*51 | High risk of Behçet Syndrome | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1199 | B*08 B*51 | HLA-B*51 | Positive The patient has the HLA antigen (HLA-B51) associated with Behcets | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | Yes
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | HLA-B*51 - Positive | HLA-B*51 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1239 | HLA-B51 POSITIVE | HLA-B*51 | The presence of the HLA-B51 genotype is associated with, but not diagnostic of, Behçet's syndrome. HLA-B51 is present in ~9% of the UK population, although has been reported to be up to 29% in other populations. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1265 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | Presence of B*51 | B*51 |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | Yes
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Partially Tested | Other | Our laboratory tests only HLA for Coeliac Desease | 2026-07-06 | 2026-07-08 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-3-2026 Rheumatoid
8-3-2026 | Rheumatoid
Distribution Date: 30/06/2026
Total distributed | 6
Total submitted | 52
Not tested | 1
Reference | DRB1*03:01, 04:07
Number acceptable | 5
Assessed | Yes
Comments
Lab Code | Rheumatoid Results for Assessment | Rheumatoid Alleles of Interest | Rheumatoid Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | No
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | No
Lab-1038 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | DRB1*03,DRB1*04 | DRB1*04 | DRB1*04=present | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1067 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | No
Lab-1079 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | DRB1*03:01,*04:07 | DRB1*04:01,*04:04,*04:05,*04:08 DRB1*10 DRB1*01:01,*01:02,*01:04 DRB1*14:06 | Absence of any susceptibility epitope carried by an HLA-DRB1 molecule. The patient has no epitope associated with a risk of developing rheumatoid arthritis. 20% of the patients with rheumatoid arthritis do not have a susceptibility epitope, so this result alone does not exclude the diagnosis. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-14 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | DRB1*03:01, DRB1*04:07 | DRB1*04:01/04:04/04:05. DRB1*01:01, DRB1*14:02 to a lesser extent | This patient does not have a HLA-DRB1 allele classically associated with Rheumatoid Arthritis (RA). | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-DRB1*03:01, 04:07 | HLA-DRB1*01:01, DRB1*01:02, DRB1*04:01, DRB1*04:04, DRB1*04:05, DRB1*04:08, DRB1*10:01, DRB1*14:02 | HLA-DRB1*01:01, DRB1*01:02, DRB1*04:01, DRB1*04:04, DRB1*04:05, DRB1*04:08, DRB1*10:01, DRB1*14:02 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-08-01 | 2026-08-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | No
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1265 | DRB1*03:01 DRB1*04:07 | DRB1*01:01/02/04 ; DRB1*10 ; DRB1*14:02 DRB1*04:01/04/05/08 | Investigation of Genetic Susceptibility to Rheumatoid Arthritis ABSENCE of a susceptibility epitope carried by a molecule HLA-DRB1. The patient does not have an epitope associated with a risk of developing rheumatoid arthritis, but 20% of patients with rheumatoid arthritis do not have one. This result alone does not exclude the diagnosis of rheumatoid arthritis (ACR 2010 criteria). (Ref : Balandraud N, Picard C, Reviron D, Landais C, Toussirot E, Roudier J. HLA-DRB1 Genotypes and the Risk of Developing Anti- Citrullinated Protein Antibody (ACPA) Positive Rheumatoid Arthritis. PLoS One (2013) 8:64108) | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | No
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Partially Tested | Other | Our laboratory tests only HLA for Coeliac Desease | 2026-07-06 | 2026-07-08 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-3-2026 Diabetes
8-3-2026 | Diabetes
Distribution Date: 30/06/2026
Total distributed | 8
Total submitted | 52
Not tested | 1
Reference | DRB1*03:01, 04:07; DQA1*03:03, 05:01; DQB1*02:01, 03:01
Number acceptable | 6
Assessed | Yes
Comments
Lab Code | Diabetes Results for Assessment | Diabetes Alleles of Interest | Diabetes Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | No
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | HLA-DRB1*03 :01 ;DQA1*05 :01 ;DQB1*02 :01 PRESENT HLA-DRB1*04 :01 ;DQA1*03 :01 ;DQB1*03 :02/04 ABSENT HLA-DRB1*04 :02 ;DQA1*03 :01 ;DQB1*03 :02/04 ABSENT HLA-DRB1*04 :04 ;DQA1*03 :01 ;DQB1*03 :02/04 ABSENT HLA-DRB1*04 :05 ;DQA1*03 :01 ;DQB1*03 :02/04 ABSENT HLA-DRB1*15 :01 ; DQA1*01 :02 ; DQB1*06 :02 ABSENT HLA-DRB1*14 :01 ; DQA1*01 :01 ; DQB1*05 :03 ABSENT HLA-DRB1*07 :01 ; DQA1*02 :01 ; DQB1*03 :03 ABSENT HLA-DRB1*04 :03 ; DQA1*03 :01 ; DQB1*03 :02 ABSENT | HLA-DRB1*03 :01 ;DQA1*05 :01 ;DQB1*02 :01 HLA-DRB1*04 :01 ;DQA1*03 :01 ;DQB1*03 :02/04 HLA-DRB1*04 :02 ;DQA1*03 :01 ;DQB1*03 :02/04 HLA-DRB1*04 :04 ;DQA1*03 :01 ;DQB1*03 :02/04 HLA-DRB1*04 :05 ;DQA1*03 :01 ;DQB1*03 :02/04 HLA-DRB1*15 :01 ; DQA1*01 :02 ; DQB1*06 :02 HLA-DRB1*14 :01 ; DQA1*01 :01 ; DQB1*05 :03 HLA-DRB1*07 :01 ; DQA1*02 :01 ; DQB1*03 :03 HLA-DRB1*04 :03 ; DQA1*03 :01 ; DQB1*03 :02 | Présence d'allèles pouvant constituer un haplotype de susceptibilité HLA DRB1*03:01/DQA1*05:01/DQB1*02:01, conférant à l'individu porteur un risque majoré de développer un diabète de type 1. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | No
Lab-1038 | HLA genotype : HLA-DRB1*03:01, DRB1*04:07 HLA-DQA1*05:01, DQA1*03:03 HLA-DQB1*02:01, DQB1*03:01 Result : HLA-DRB1*03:01, DQA1*05:01, DQB1*02:01 : presence HLA-DRB1*04:01/04:02/04:04/04:05, DQA1*03, DQB1*03:02 : absence HLA-DRB1*04:01/04:02/04:04/04:05, DQA1*03, DQB1*03:04 : absence | HLA-DRB1*03:01, DQA1*05:01, DQB1*02:01 HLA-DRB1*04:01/04:02/04:04/04:05, DQA1*03, DQB1*03:02 HLA-DRB1*04:01/04:02/04:04/04:05, DQA1*03, DQB1*03:04 | This individual carries HLA-DRB1*03:01, DQA1*05:01, DQB1*02:01 (DR3,DQ2) which is associated with type 1 diabetes. The DR3,DQ2 and DR4,DQ8 haplotypes are found in 95% of type 1 diabetes patients whereas the HLA-DR3 and DR4 antigens are found 40% of the Caucasian population. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | No
Lab-1079 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | DRB1*03:01,*04:07 DQA1*03:03,*05:01 DQB1*02:01,*03:01 | DRB1*03:01;DQA1*05:01;DQB1*02:01 DRB1*04:01;DQA1*03:01;DQB1*03:02/04 DRB1*04:02;DQA1*03:01;DQB1*03:02/04 DRB1*04:04;DQA1*03:01;DQB1*03:02/04 DRB1*04:05;DQA1*03:01;DQB1*03:02/04 | Presence of alleles that may constitute a susceptibility haplotype HLA-DRB1*03:01/DQA1*05:01/DQB1*02:01, giving the carrier an increased risk of developing type I diabetes, with an odds ratio of 3,64. This result is not a diagnostic criterion for the disease. | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-14 | Yes
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | DQB1*02:01, DQB1*03:01 | DQB1*03:02, DQB1*02 | This patient has a HLA genotype (HLA-DQB1*02:01) associated with Insulin-Dependent Diabetes Mellitus (IDDM). | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-DRB1*03:01, 04:07; DQA1*03:03, 05:01; DQB1*02:01, 03:01 | HLA-DRB1*03:01/DQA1*05:01/DQB1*02:01; HLA-DRB1*0405/DQA1*03:01/03/DQB1*03:02; HLA-DRB1*04:01/DQA1*03:01/03/DQB1*03:02; HLA-DRB1*04:02/DQA1*03:01/DQB1*03:02 | HLA-DRB1*03:01/DQA1*05:01/DQB1*02:01 positive | Acceptable |  | Yes | Fully Tested |  |  | 2026-08-01 | 2026-08-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | No
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | HLA-DRB1*03:01 HLA-DRB1*04:07 | HLA-DRB1* |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | Yes
Lab-1239 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1265 | DQB1*02:01 ; DQA1*05:01 ; DRB1*03:01 DQB1*03:01 ; DQA1*03:02 ; DRB1*04:07 | DQB1*02:01/DQA1*05:01/DRB1*03:01 ; DQB1*03:02/DQA1*03/DRB1*04:01/02/04/05 | Investigation of susceptibility to type I diabetes PRESENCE of haplotype HLA-DRB1*03:01/DQB1*02:01/DQA1*05:01(DR3-DQ2) associated with an increased risk of developing type 1 diabetes. These results do not constitute a diagnostic criterion for the disease. No specific preventive measures, nor family or antenatal research, are required in carriers of these predisposition genes. Absence of haplotype HLA-DRB1*04:xx/DQA1*03/DQB1*03:02(DR4-DQ8) (Ref: Noble JA. Immunogenetics of type 1 diabetes: A comprehensive review. J Autoimmun. nov 2015;64:101-12. Zhao LP et al nine residues in HLA-DQ molecules determine with susceptibility and resistance to type 1 diabetes among young children in Sweden. Sci Rep . 2021 Apr 23;11(1):8821) | Unacceptable | Reported DQA1*03:02 instead of DQA1*03:03 | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | Yes
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | No
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Partially Tested | Other | Our laboratory tests only HLA for Coeliac Desease | 2026-07-06 | 2026-07-08 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes

## 8-3-2026 Psoriasis
8-3-2026 | Psoriasis
Distribution Date: 30/06/2026
Total distributed | 6
Total submitted | 52
Not tested | 1
Reference | C*07:01, 15:02
Number acceptable | 5
Assessed | Yes
Comments
Lab Code | Psoriasis Results for Assessment | Psoriasis Alleles of Interest | Psoriasis Interpretive Comments | Assessment | Assessment notes | isSubmitted | Tested | Reason for not testing | Confirm Reason | Date Received | Date Tested | Registered
Lab-1003 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-08 | No
Lab-1005 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1008 | C*07:01/06/16/18/36++, *15:02/04/09/13/17++; | HLA-C*06 | This patient does not have the HLA-C*06 allele group associated with psoriasis | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | Yes
Lab-1011 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1020 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1024 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1025 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-13 | No
Lab-1026 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | No
Lab-1027 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-29 | 2026-07-02 | No
Lab-1031 |  | Not tested |  |  |  | Yes | Fully Tested |  |  | 2026-07-07 | 2026-07-16 | No
Lab-1033 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-06 | 2026-07-08 | No
Lab-1038 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1039 | [5 empty] | Yes | Fully Tested |  |  | 2026-06-30 | 2026-07-06 | No
Lab-1057 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | No
Lab-1061 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-09 | 2026-07-13 | No
Lab-1065 | C*06 negative | C*06 |  | Acceptable |  | Yes | Fully Tested |  |  |  |  | Yes
Lab-1067 | N/A |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1077 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-06 | No
Lab-1079 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1080 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1085 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1086 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1090 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-14 | No
Lab-1104 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1111 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1118 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1120 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1132 | Not performed. Not enrolled in the program. |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-03 | No
Lab-1134 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-07 | No
Lab-1139 | B*08, B*51 C*07, C*15 | C*06 or B*27 | This patient does not have the HLA alleles associated with psoriasis | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-03 | Yes
Lab-1145 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-16 | No
Lab-1146 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1150 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-21 | No
Lab-1151 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-10 | 2026-07-20 | No
Lab-1154 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-02 | No
Lab-1176 | HLA-C*07:01, 15:02 | HLA-C*06 | HLA-C*06 negative | Acceptable |  | Yes | Fully Tested |  |  | 2026-08-01 | 2026-08-03 | Yes
Lab-1186 | [5 empty] | Yes | Fully Tested |  |  |  |  | No
Lab-1199 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-15 | No
Lab-1205 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1206 | Number of HLA-C*06:02 alleles: 0. SNP genotypes are rs10484554: CC, rs3130457: CT, rs6904246: GG, rs7745906: GG. |  |  | Acceptable |  | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-07 | Yes
Lab-1207 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1208 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-09 | No
Lab-1209 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-17 | No
Lab-1215 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1239 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-08 | No
Lab-1265 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-08 | No
Lab-1268 | [5 empty] | No | [5 empty] | No
Lab-1269 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-22 | No
Lab-1272 | NT |  |  |  |  | Yes | Fully Tested |  |  | 2026-07-06 | 2026-08-04 | No
Lab-1284 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-08 | No
Lab-1291 | [5 empty] | Yes | Partially Tested | Other | Our laboratory tests only HLA for Coeliac Desease | 2026-07-06 | 2026-07-08 | No
Lab-1295 | // | // | // |  |  | Yes | Fully Tested |  |  | 2026-07-03 | 2026-07-10 | No
Lab-1296 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-01 | 2026-07-06 | No
Lab-1304 | [5 empty] | Yes | Fully Tested |  |  | 2026-07-02 | 2026-07-06 | No
Lab-1306 | NT |  |  | Not Tested |  | Yes | Not Tested | Samples not received | Shipment refused by customs |  |  | Yes